## 7.2 Help text
General : If the value of a field is unknown and the field is not mandatory please leave blank. Only fields 2.1 a,b,c default to 'unknown' as this is a requirement for electronic exchange. For the rest of fields please leave blank if unknown. The form should be used to report Incidents and Serious Incidents under AIMDD/MDD/IVDD and MDR/IVDR respectively’
|  |  | Mandatory properties |  |  |  |  | Comments
| Section | Help text | Combined initial & final | Initial | Follow Up | Final (Reportable incident) | Final (Non-reportable incident)
1. Administrative information
1.1 | Corresponding competent authority
a | Name of receiving national competent authority (NCA) | NCA to which the report is being sent List is available here: https://ec.europa.eu/growth/sectors/medical-devices/contacts_en under Vigilance contact points. | Y | Y | Y | Y | Y
b | EUDAMED Number of NCA | Unique Eudamed number of NCA (could be auto filled/selected once Eudamed available) | N | N | N | N | N | Will be mandatory as soon as available in EUDAMED
c | Reference number assigned by NCA | The reference number of the NCA for this incident (if known- i.e. follow up and finals). If not known please add 'not known' | N | N | Y | Y | Y
d | Reference number assigned by EUDAMED for this incident | The reference number assigned by Eudamed for this incident (after upload/entering incident into Eudamed) | N | N | N | N | N | Will be mandatory as soon as available in EUDAMED
1.2 | Date, type, and classification of incident report
a | Date of submission | Date when you submit the report - . The predefined date format is: YYYY-MM-DD | Y | Y | Y | Y | Y
b | Date of incident | Date when incident happened - if incident date is unknown please add the dates between when you think the incident occurred. If this incident is a result of a literature report please enter a date range similar to the range of the report. In case the specific date is known you only have to enter it once into the first field. The same date is automatically entered into the second field. In case of a time-span you need to adjust the second field accordingly. | Y | Y | Y | Y | Y
c | Manufacturer awareness date | Date when company employee was made aware of the incident | Y | Y | Y | Y | Y
d | Type of Report | Initial Report - Choose this for your first report on a particular incident. It may be initiated under the Vigilance system, or may stem from a User Incident Report notified to you by the relevant CA. Follow-up report - Choose this to provide additional/interim information during your investigation. You can submit more than one follow-up report for each incident. Note this option is only available once you have submitted an initial report. Combined initial & final - Choose this if you have the details of the initial and final report within the initial timeframe for reporting. Note that this option is not available once you have submitted an initial report. Final (Reportable incident)- This is your formal statement of the outcome of your investigation, including any actions proposed or taken. It is recognised that a further ‘Final’ report may be necessary in circumstances where additional information only becomes available at a later stage. Note this option is only available once you have submitted an initial report. Applicable for both vigilance reports submitted in compliance with MDD/IVDD (30 days) and for MDR/IVDR (15 days) Final (Non-reportable incident)- This is to be used if the event reported in the initial vigilance report is found after analysis not to fulfil the reporting criteria. By selecting this the only mandatory field is the 4.2B and no other fields in Section 4 It is possible to send an update at any stage to each type of reports except for the initial report e.g. update to the final report. | Y | Y | Y | Y | Y
e | In case of initial and follow-up reports, please indicate the expected date of next report | The date by which you expect to be able to submit your next report on this event. This may be either a Follow-Up or Final report. | N | Y | Y | N | N
f | Classification of incident | Please classify the incident according with the following definitions: Serious public health threat: An event which could result in imminent risk of death, serious deterioration in a person's state of health, or serious illness, that may require prompt remedial action, and that may cause significant morbidity or mortality in humans, or that is unusual or unexpected for the given place and time Death: The event led to the death of a patient, user or other person. Unanticipated serious deterioration in state of health: An unanticipated serious deterioration in state of health is one where the event was not considered in a risk analysis, and where it led to a serious deterioration in the state of health of a patient, user or other person. All other reportable incidents: These are events which did not involve a death and which were not unanticipated but which led, or might have led, to a serious deterioration in the state of health of a patient, user or other person. | Y | Y | Y | Y | N
1.3 | Submitter information
1.3.1 a | Submitter of report | Who is submitting the report | Y | Y | Y | Y | Y
b | Manufacturer's reference number for this incident | The reference number assigned by the manufacturer | Y | Y | Y | Y | Y
c | If this incident involves multiple devices from the same manufacturer, please list the respective reference numbers of the other MIR forms you have submitted | If other devices were involved in the incident, you must send a separate MIR form for each device. List the reference numbers of the NCA, EUDAMED and manufacturer here. (As each suspected device will have its own report as opposed to section 2.6 where only accessories/associated to the suspected devices are listed ) Please use a semi colon to separate multiple values 1.3.c, 2.6: If you are certain of the device that caused the incident, then this device is what should be reported on the form – List any associated accessories in 2.6a which could be from a different manufacturer if known and 2.6b for any other devices (which could also be from a different manufacturer if known) If unsure then submit separate MIR reports on all other devices from the same manufacturer that could have played a role and fill out 1.3.1c. List any associated accessories in 2.6a which could be from a different manufacturer if known and 2.6 b for any other devices (which could be from a different manufacturer if known) and this will indicate to the CA that those other manufacturers should also submit a MIR. | N | N | N | N | N
d | If this incident is covered under an FSCA, please provide the relevant numbers: | NCA's local FSCA reference - The NCA reference number assigned to the FSCA that this incident covers EUDAMED's FSCA reference number Manufacturer's FSCA reference number - If the FSCA sent to the NCA contained multiple issues, the NCA may assign a unique number to each issue. If this incident in this report is related to one of those issues, list the unique number assigned by the NCA for that issue in the field "NCA's local FSCA reference" Please use a semi colon to separate multiple values | N | N | N | N | N
e | Periodic Summary Report (PSR) ID | Please quote the unique PSR-ID for the incident if it is reportable under PSR. If under PSR then you must fill this out (The PSR task force is working up a new methodology for PSR submission . This field is here as a placeholder to facilitate this new method of submission.) Please use a semi colon to separate multiple values | N | N | N | N | N
f | If the incident occurred within a PMCF/ PMPF investigation; please provide the Eudamed ID of that PMCF/PMPF investigation | If the incident occurred during a PMCF/PMPF investigation then you must fill this out. The word 'studies' is generally used for PMCF/PMPF however the MDR refers to this as 'investigation' | N | N | N | N | N
1.3.2 | Manufacturer information
a | Manufacturer Organisation name | The name of the Manufacturer for the device involved in this adverse incident | Y | Y | Y | Y | Y
b | Single Registration Number (SRN) | SRN is the unique identifier which will be the unique identifier of actors in the future Eudamed. When an SRN is available, an SRN field will be completed and will pre-populate the manufacturer details, including the Manufacturer's name. | Y | Y | Y | Y | Y | When available will be mandatory. The full details will be auto filled once SRN is available
c | Contact's first name | First name of the manufacturer contact person | Y | Y | Y | Y | Y | Information in c,d, e, f will be auto populated when in Eudamed. These fields will be editable so can be overwritten/updated. If there is an Authorised Representative, the AR will be the preliminary contact for CA when manufacturer is outside Europe. Please also include a manufacturer contact if possible (as well as the Authorised Representative.)
d | Contact's last name | Last name of the manufacturer contact person | Y | Y | Y | Y | Y | Information in c,d, e, f will be auto populated when in Eudamed. These fields will be editable so can be overwritten/updated. If there is an Authorised Representative, the AR will be the preliminary contact for CA when manufacturer is outside Europe. Please also include a manufacturer contact if possible (as well as the Authorised Representative.)
e | E-mail | E-mail of the manufacturer contact person . | Y | Y | Y | Y | Y | Information in c,d, e, f will be auto populated when in Eudamed. These fields will be editable so can be overwritten/updated. If there is an Authorised Representative, the AR will be the preliminary contact for CA when manufacturer is outside Europe. Please also include a manufacturer contact if possible (as well as the Authorised Representative.)
f | Phone | Telephone number of the manufacturer contact person | Y | Y | Y | Y | Y | Information in c,d, e, f will be auto populated when in Eudamed. These fields will be editable so can be overwritten/updated. If there is an Authorised Representative, the AR will be the preliminary contact for CA when manufacturer is outside Europe. Please also include a manufacturer contact if possible (as well as the Authorised Representative.)
g | Country | The country where the manufacturer is located. (Note for development- drop down with ISO code and it should be worldwide) | Y | Y | Y | Y | Y
h | Street | The street of the manufacturer | Y | Y | Y | Y | Y | One of either h, i, j, k
i | Street number | The street number of the manufacturer | N | N | N | N | N | One of either h, i, j, k
j | Address complement | The address where the Manufacturer is located. E.g. building name | N | N | N | N | N | One of either h, i, j, k
k | P.O Box | The P.O box of the manufacturer | N | N | N | N | N | One of either h, i, j, k
l | City name | The name of the city where the manufacturer is located. | Y | Y | Y | Y | Y
m | Postcode | The postal or zip code where the manufacturer is located | Y | Y | Y | Y | Y
1.3.3 | Authorized representative information
a | Authorised representative Organisation name | The name of the Authorised representative for the device involved in this adverse incident | Y | Y | Y | Y | Y | Mandatory if manufacturer is non EEA, CH, TR
b | Single Registration Number (SRN) | SRN is the unique identifier which will be the unique identifier of actors in the future Eudamed. When an SRN is available, an SRN field will be completed and will pre-populate the Authorised representative details, including the Authorised representative's name. | Y | Y | Y | Y | Y | When available will be mandatory. The full details will be auto filled once SRN is available
c | Contact's first name | First name of the Authorised representative contact person | Y | Y | Y | Y | Y
d | Contact's last name | Last name of the Authorised representative contact person | Y | Y | Y | Y | Y
e | E-mail | E-mail of the Authorised representative contact person . | Y | Y | Y | Y | Y
f | Phone | Telephone number of the Authorised representative contact person | Y | Y | Y | Y | Y
g | Country | The country where the Authorised representative is located. | Y | Y | Y | Y | Y
h | Street | The street of the Authorised representative | Y | Y | Y | Y | Y | One of either h, i, j, k
i | Street number | The street number of the manufacturer Autorised representative | N | N | N | N | N | One of either h, i, j, k
j | Address complement | The address where the Authorised representative is located. E.g. building name | N | N | N | N | N | One of either h, i, j, k
k | P.O Box | The P.O box of the Authorised representative | N | N | N | N | N | One of either h, i, j, k
l | City name | The name of the city where the Authorised representative is located. | Y | Y | Y | Y | Y
m | Postcode | The postal or zip code where the Authorised representative is located | Y | Y | Y | Y | Y
1.3.4 | Submitters details if not manufacturer or authorized representative
a | Registered commercial name of company | The name of the company (the legal entity) for this adverse incident | Y | Y | Y | Y | Y | Only mandatory if Submitter of reporter is 'Other'
b | Contact's first name | First name of the person to contact about the incident | Y | Y | Y | Y | Y | Only mandatory if Submitter of reporter is 'Other'
c | Contact's last name | Last name of the person to contact about the incident | Y | Y | Y | Y | Y | Only mandatory if Submitter of reporter is 'Other'
d | E-mail | The email address for the submitter | Y | Y | Y | Y | Y | Only mandatory if Submitter of reporter is 'Other'
e | Phone | The telephone number for the company | Y | Y | Y | Y | Y | Only mandatory if Submitter of reporter is 'Other'
f | Country | The country where the company is located. (Note for development- drop down with ISO code and it should be worldwide) | Y | Y | Y | Y | Y | Only mandatory if Submitter of reporter is 'Other'
g | Street | The street of the submitter | Y | Y | Y | Y | Y | One of either g, h, i, j
h | Street number | The number of the manufacturer on the street | Y | Y | Y | Y | Y | One of either g, h, i, j
i | Address complement | The address where the company is located- e.g. building name | N | N | N | N | N | One of either g, h, i, j
j | P.O Box | The P.O box of the submitter | N | N | N | N | N | One of either g, h, i, j
k | City name | The name of the city where the company is located. | N | N | N | N | N
l | Postcode | The postal or zip code where the company is located | Y | Y | Y | Y | Y
2. Medical Device Information
2.1 | Unique Device Identification (UDI)
a | UDI device identifier/Eudamed ID | The UDI Device Identifier (DI). In cases where devices are supplied within a system or procedure pack, if an individual device is being reported on, use the UDI-DI for the device and record the system or procedure pack information in 2.6 as an associated device. If the system or procedure pack is being reported on, then use the UDI information for the system or procedure pack. In the event of an incident with a legacy device (regardless of this device situation on the market) such device has to be registered in Eudamed. In that case, the legacy device will have be assigned a Eudamed ID instead of a UDI Device Identifier (DI). | Y | Y | Y | Y | Y | default value to be ‘Unknown’ and for MDD/IVD risk class field is not mandatory but for MDR/IVDR risk class field is mandatory. Form will show ‘unknown’ and it is up to the manufacturer to overwrite (or auto populate in Eudamed). Having the field prepopulated with 'unknown' ensures consistency when the actual UDI is unknown. There is potential for error if it is typed in and the auto processing of the report will fail.
b | UDI production identifier | The UDI Production Identifier (PI) - if unknown at time of submission please leave as 'unknown' and fill out on follow up or final. If two devices have failed in the same way you must fill out two MIR forms. Please do not add multiple UDI PI's | Y | Y | Y | Y | Y | default value to be ‘Unknown’ and for MDD/IVD risk class field is not mandatory but for MDR/IVDR risk class field is mandatory. Form will show ‘unknown’ and it is up to the manufacturer to overwrite (or auto populate in Eudamed).Having the field prepopulated with 'unknown' ensures consistency when the actual UDI is unknown. There is potential for error if it is typed in and the auto processing of the report will fail.
c | Basic UDI-DI/Eudamed DI | The Basic UDI-DI is the main key in the database and relevant documentation (e.g. certificates, declaration of conformity, technical documentation and summary of safety and clinical performance) to connect devices with same intended purpose, risk class and essential design and manufacturing characteristics. It is independent/separate from the packaging/labelling of the device and it does not appear on any trade item. Any Basic UDI-DI shall identify the devices (group) covered by that Basic UDI-DI in a unique manner. Basic UDI-DI is only required for Regulation-compliant product reporting – see MDCG approved guideline MDCG 2019 -5) It auto populates when entering UDI-DI in Section 2.1.a. In the event of an incident with a legacy device (regardless of this device situation on the market) such device has to be registered in Eudamed. In that case, the legacy device will have be assigned a Eudamed ID instead of a Basic UDI-DI | Y | Y | Y | Y | Y | default value to be ‘Unknown’ and for MDD/IVD risk class field is not mandatory but for MDR/IVDR risk class field is mandatory. Form will show ‘unknown’ and it is up to the manufacturer to overwrite (or auto populate in Eudamed).Having the field prepopulated with 'unknown' ensures consistency when the actual UDI is unknown. There is potential for error if it is typed in and the auto processing of the report will fail.
d | Unit of use UDI-DI | Identifier assigned to an individual medial device when a UDI-DI is not labelled on the individual device at the level of its unit of use, for example in the event of several units of the same device being packaged together. | N | N | N | N | N
2.2 | Categorisation of device
a | Medical Device terminology used | Select the CND code of your product. Select the code system that your organisation uses for legacy devices (devices with valid Directive certificate placed on the market after date of application of the Regulation). | N | N | N | N | N
b | Medical device nomenclature code | Code from the terminology above | N | N | N | N | N
2.3 | Description of device and commercial information | (IMDRF definitions from the published 'Common Data Elements for Medical Device Identification' document are used here- indicated with IMDRF after the definition. (http://www.imdrf.org/docs/imdrf/final/technical/imdrf-tech-rps-common-data-elements.pdf)
a | Medical device name (Brand/Trade/Proprietary or Common name) | A name used to assist in the identification of the regulated medical device (IMDRF) - If unknown , put 'unknown' and update when you receive the information. | Y | Y | Y | Y | Y
b | Nomenclature text/Description of the device and its intended use | Nomenclature text: can be the same as the text defining the code in 2.2b Or Description of the device and its intended use: Please describe the intended use of the device as foreseen in the IFU. | Y | Y | Y | Y | Y
c | Model | The value used to represent one medical device or a family of medical devices to group many variations that have shared characteristics (IMDRF). | N | N | N | N | N
d | Catalogue/Reference number | The value given by the Regulated Entity to identify the specific medical device as it relates to its form/fit, function and process (i.e., manufacturing processes requiring differentiation for distribution control (e.g., sterilization, component material, reprocessing, etc.) (IMDRF) | N | N | N | N | N
e | Serial number(s) | The serial number(s) for the medical device to which this report refers | N | N | N | N | N
f | Lot/batch number(s) | The Lot/batch number(s) number for the medical device to which this report refers | N | N | N | N | N
g | Software version | The value given by the applicant to identify a specific revision of the software (including SaMD) (IMDRF) | N | N | N | N | N
h | Firmware version | The value given by the applicant to identify a specific revision of the firmware (including SaMD) (IMDRF) | N | N | N | N | N
i | Device manufacturing date | The date of manufacture for the medical device to which this report refers | N | N | N | N | N
j | Device expiry date | The expiry date for the medical device to which this report refers | N | N | N | N | N
k | Date when device was implanted | If available, the implant date for the medical device to which this report refers- if exact date is unknown please add timeframe | N | N | N | N | N
l | Date when device was explanted | If avaialble, the explant date for the medical device to which this report refers- if exact date is unknown please add timeframe | N | N | N | N | N
m | If precise implant/explant dates are unknown, provide the duration of implantation | How long was the implant in the patient? (months/years) Please provide best estimate. If day is unknown, month and year are acceptable. If month and day are unknown, year is acceptable. | N | N | N | N | N
n | Implant facility | The healthcare facility where the device was implanted | N | N | N | N | N
o | Explant facility | The healthcare facility where the device was explanted | N | N | N | N | N
p | Notified body (NB) ID number(s) (if applicable ) | The most current identification number for the Notified Body that granted the CE mark for this device. If two Notified Bodies are involved, please enter both values in the field provided | N | N | N | N | N | Depends on Risk Class in 2.4 - see sheet labelled 'Risk Class and NB properties'
p | Notified body (NB) certificate number(s) of device (if applicable) | The certificate number from the notified body(ies) above. If two Notified Bodies, please enter both values in the fields provided. (For example, kit containing multiple components) . (For example, a kit containing multiple components.) | N | N | N | N | N | Depends on Risk Class in 2.4 - see sheet labelled 'Risk Class and NB properties'
q | Please indicate the date of one of the following: | What was the date of either: the first Declaration of Conformity; or the device first CE marked, or first placed on the market and/or put into service or if software, the first date available for download. For CE marked devices: the date of the first Declaration of Conformity or the date the CE Marking was affixed (these dates are usually considered to coincide, but may not). For Custom-Made devices: the date the device was first placed on the market and/or put into service; For In-House manufactured devices: the date the device was first made available for use and/or put into service. Please indicate to which option the date refers. Please note: single choice | N | N | N | N | N
2.4 | Risk class of device when placed on market | Select which regulation the device was conforming to when it was placed on the market, irrespective of when the actual incident occurred
a | This device has been placed on the market before the implementation of the MDD/AIMDD/IVDD | Select if the device was placed on the market before the implementation of the MDD (93/42/EEC)/AIMDD (90/385/EEC)/IVDD (98/79/EC) It is possible that on the market there are devices which were placed on the market before June 1993 or active implants implanted before 1990 or an IVD before 1998. | Y | Y | Y | Y | Y | One of either a,b or c
b | MDD/AIMDD/IVDD Class | Indicate the relevant Risk Class for the medical device to which this report refers. If unknown please select risk class to the best of your knowledge and update on follow up/final Auto populates when entering UDI-DI in Section 2.1.a. | Y | Y | Y | Y | Y | One of either a,b or c
b | MDR/IVDR Class | MDR and IVDR fields set up as according to proposal in Eudamed The information will be attached to the UDI (Basic UDI-DI) in the UDI database, therefore it will come automatically with the Basic UDI-DI (except for custom-made device which will not have a Basic UDI-DI). If unknown please select risk class to the best of your knowledge and update on follow up/final. Auto populates when entering UDI-DI in Section 2.1.a. | Y | Y | Y | Y | Y | One of either a,b or c
2.5a | Market distribution of device (region/ country) (according to the best knowledge of the manufacturer) | Indicate which countries the medical device has been distributed to- please fill to the best of your knowledge Auto populates when entering UDI-DI in Section 2.1.a. | Y | N | N | Y | N
2.6 | Use of accessories, associated devices or other devices
a | Relevant accessories used with the device being reported on, if applicable | Accessories enable the medical device to be used for its intended purpose and may or may not be from the same manufacturer as the device being reported on. Provide as much information as possible, including manufacturer names if different. This field may be used to alert the relevant CA of the need for a separate MIR form from the other manufacturer(s) involved. 1.3.c, 2.6: If you are certain of the device that caused the incident, then this device is what should be reported on the form – List any associated accessories in 2.6a which could be from a different manufacturer if known and 2.6b for any other devices (which could also be from a different manufacturer if known) If unsure then submit separate MIR reports on all other devices from the same manufacturer that could have played a role and fill out 1.3.1c. List any associated accessories in 2.6a which could be from a different manufacturer if known and 2.6 b for any other devices (which could be from a different manufacturer if known) and this will indicate to the CA that those other manufacturers should also submit a MIR. | N | N | N | N | N
b | Relevant associated devices used with the device being reported on, if applicable | Associated devices are those used together/in combination and may or may not be from the same manufacturer as the device being reported on. An example is the use of a CT scanner used in combination with contrast medium; contrast injectors; tubing, syringes etc. Provide as much information as possible, including manufacturer names if different. This field may be used to alert the relevant CA of the need for a separate MIR form from the other manufacturer(s) involved. | N | N | N | N | N
3.Incident information derived from health care professional and/or user
3.1 | Nature of incident
a | Provide a comprehensive description of the incident, including (1) what went wrong with the device (if applicable) and (2) a description of the health effects (if applicable), i.e. clinical signs, symptoms, conditions as well as the overall health impact | Describe the incident/event including any relevant information that might impact the understanding or evaluation of the adverse event. This should cover the first observable event or possibly a later secondary but more serious event, or maybe both. Describe the condition/outcome of a patient after the event, including the degree of wellness and the need for continuing care, medication, support, counselling, or education. | Y | Y | Y | Y | Y
3.2 | Medical device problem information
a | IMDRF Medical device problem codes (annex A) | Please enter the most relevant Level 3 observation on the medical device problem as Choice 1. You may chose up to 5 other different codes that describe the medical device problem. In case you can't find a level 3 observation, but a suitable level 2 code, then please use this code as Choice 1 and explain briefly in the text box why no level 3 code was chosen. You can propose a new IMDRF code/term. If no code on level 2 can be chosen, select one from level 1 and propose a new IMDRF code/term for levels 2 and 3. The aim is always to code to the most appropriate level. If no code (on any hierarchy level) can be found briefly explain why. (This is a way to propose new IMDRF terms which could be incorporated in the nomenclature during a maintenance session) For some codes in Annex A there is no level 3. In those cases, level 1 or 2 would be acceptable without justification. Note Choice 1 is mandatory after the transitional period. The other codes are optional. | Y | Y | Y | Y | Y | Mandatory once transition period is over
b | Number of patients involved | The number of patients involved . Please note that Patient could also mean user or other third person and the initial reporter may be a family members or other third person. If issue occurred on External Quality Assessment Samples and no patient was involved, should this field be “0”. | N | N | N | N | N
c | What is the current location of the device | The current location of the device involved in this event- please describe the location if none of the provided options are applicable. Please Note- single choice | Y | Y | Y | Y | Y
d | Operator of device at the time of the incident | Indicate who was operating the device at the time of the event: healthcare professional, patient, Other: please describe the operator if none of the provided options are applicable. Please Note- single choice | N | N | N | N | N
e | Usage of device | Indicate the usage of the device at the time the incident occurred . Please describe the usage if none of the provided options are applicable select Other and describer. Compassionate / humanitarian use For reagents/test strips select “initial use” as the assay itself will not be re-used (i.e. another drop of reagent or another reaction support is going to be used for the next test/sample). For instruments/analysers, select “Reuse of a reusable medical device” as the analyser is going to be used several years to perform testing. Fully refurbished: see NB-MED/2.1/Rec5 guidance | N | N | N | N | N
f | Remedial actions taken by healthcare facility, patient or user subsequent to the incident | Describe any action taken by the health practitioner, healthcare facility, patient or user to avoid a further occurrence of this problem. This may include explant of an implanted device; the recall of patients for further testing; or the provision of changed or improved advice on the use of the device or the quarantine of possibly affected devices / device involved in this incident. | N | N | N | N | N
3.3 | PATIENT INFORMATION
a | IMDRF health effect codes | Mandatory once transition period is over for using IMDRF 'Clinical signs, symptoms, and conditions codes' (Annex E) and IMDRF 'Health impact' codes (Annex F)". You can choose up to 6 codes to describe: 1) clinical signs, symptoms, conditions and 2) health impact. Please enter the most relevant lowest level observation as Choice 1. You may chose up to 5 other different codes that describe the health effect In case you can't find a level 3 observation, but a suitable level 2 code, then please use this code as Choice 1 and explain briefly in the text box why no level 3 code was chosen. You can propose a new IMDRF code/term. If no code on level 2 can be chosen, select one from level 1 and propose a new IMDRF code/term for levels 2 and 3. The aim is always to code to the most appropriate level. If no code (on any hierarchy level) can be found briefly explain why. (This is a way to propose new IMDRF terms which could be incorporated in the nomenclature during a maintenance session) For some codes there is no level 3. In those cases, level 1 or 2 would be acceptable without justification. Or standardized reply possible? IMDRF Adverse Event Annex E, F (Health Effects) codes will be published in 2019. The transition for their implementation in Europe is 12 months after their publication on the IMDRF website. | N | N | N | N | N
b | Age of the patient at the time of incident, if applicable | The age of the patient is in years or months or days. Months & days should only be used for patients under the age of 1. For all others please use years and months | N | N | N | N | N
c | Gender | The gender of the patient involved in this event | N | N | N | N | N
d | Body Weight (Kg) | The patient's weight in kilograms (Please leave blank if unknown) | N | N | N | N | N
e | List any of the patient's prior health condition or medication that may be relevant to this incident | List any health condition that may be relevant to the incident or medication taken by the patient either regularly or at the specific day/period when the incident occurred. Exclude medication taken as a consequence of the incident, i.e. to treat the adverse event | N | N | N | N | N
3.4 | INITIAL REPORTER (can be healthcare professional of facility, patient, lay user)
a | Role of initial reporter | Please select the role of the person that initially reported the incident. This can be for example: healthcare professional of facility, patient, user etc | Y | Y | Y | Y | Y
b | Name of the health care facility where incident occurred | The name of the healthcare facility where this event occurred. If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
c | Healthcare facility report Number | The Report reference number used by the healthcare facility concerned when they reported the incident. | N | N | N | N | N
d | Contact's first name | The name of the contact at the healthcare facility where this event occurred. If this incident did not occur in a health care facility you don’t need to fill out this field. | N | N | N | N | N
e | Contact's last name | The name of the contact at the healthcare facility where this event occurred. If this incident did not occur in a health care facility you don’t need to fill out this field. | N | N | N | N | N
f | Email | The telephone number for the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
g | Phone | The email address for the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
h | Country | The country of the initial reporter where this event occurred. If this even did not occur in a health care facility put the origin of the country of the report. | Y | Y | Y | Y | Y
i | Street | The street of the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
j | Street number | The street number of the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
k | Address complement | The address of the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
l | P.O Box | The P.O box of the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
m | City name | The name of the city initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
n | Postcode | The postal or zip code of the initial reporter If this incident did not occur in a health care facility you don’t need to fill out this field | N | N | N | N | N
4. Manufacturer analysis
4.1 | Manufacturer’s preliminary comments
a | For initial and follow-up reports: preliminary results and conclusions of manufacturer’s investigation | Details of any preliminary analysis you can provide | N | Y | Y | N | N
b | Initial actions (corrective and/or preventive) implemented by the manufacturer | Details of initial corrective and/or preventive actions implemented by the manufacturer. Please add n/a if not applicable yet | N | N | Y | N | N
c | What further investigations do you intend in view of reaching final conclusions? | Provide details of any further investigations you plan to undertake to reach the root cause. Please add n/a if not applicable yet | N | N | Y | N | N
4.2 | Cause investigation and conclusion
a | For Final (Reportable incident)- Description of the manufacturer’s evaluation concerning possible root causes/causative factors and conclusion | A report of your analysis of the device involved in this event. If no analysis or investigation of the device could be carried out, for example if the device was not returned, then indicate the most probable root cause analysis. Please be clear whether you are describing the root cause or most probable root cause. (Standard practice is for CA's to routinely ask for most probable root cause when it is unclear. To avoid emails going back and for the we ask for it here.) In many cases, evaluation of the actual device is not feasible. Instead the investigation may focus on "surrogate devices", i.e. devices belong to the same batch or lot as the device, to another lot/batch. Testing of model variants may also be indicated in specific cases. Please describe the devices investigated and, if applicable, tested-- Other investigative means than testing may include interviews with the user/operator of the device, control of production records etc. Your investigations and conclusions should be explained and their credibility and plausibility justified.) If you refer to an existing CAPA please provide the reference or identification number of that CAPA | Y | N | N | Y | N
b | For Final (Non-reportable incident)- Fill out rationale for why this is considered not reportable | If the incident was deemed not reportable in the first place please indicate here why. You do not have to fill out any other additional fields e.g. Two devices are involved and one of them was proved not to contribute to the (serious) incident as a result of the investigation. | N | N | N | N | Y
c | Is root cause confirmed | Please indicate if there is evidence of a causal link to the medical device problem | Y | N | N | Y | N
d | Has the risks assessment been reviewed? | Indicate if the risk assessment has been reviewed after this incident and provide a rationale why the risk assessment is still adequate or why no review is required | Y | N | N | Y | N
e | IMDRF Cause investigation terms and codes | Please enter the most relevant or most likely lowest level Cause investigation codes as Choice 1 in each section. In case you can't find a level 3 observation, but a suitable level 2 code, then please use this code as Choice 1 and explain briefly in the text box why no level 3 code was chosen. You can propose a new IMDRF code/term. If no code on level 2 can be chosen, select one from level 1 and propose a new IMDRF code/term for levels 2 and 3. The aim is always to code to the most appropriate level. If no code (on any hierarchy level) can be found briefly explain why. (This is a way to propose new IMDRF terms which could be incorporated in the nomenclature during a maintenance session) For some codes there is no level 3. In those cases, level 1 or 2 would be acceptable without justification Note Choice 1 is mandatory after the transitional period. The other codes are optional. | Y | N | N | Y | N
f | IMDRF Component codes | Please enter the most relevant lowest level component code as Choice 1 in each section. You may chose up to 5 other different codes that describes the component In case you can't find a level 3 observation, but a suitable level 2 code, then please use this code as Choice 1 and explain briefly in the text box why no level 3 code was chosen. You can propose a new IMDRF code/term. If no code on level 2 can be chosen, select one from level 1 and propose a new IMDRF code/term for levels 2 and 3. The aim is always to code to the most appropriate level. If no code (on any hierarchy level) can be found briefly explain why. (This is a way to propose new IMDRF terms which could be incorporated in the nomenclature during a maintenance session) For some codes there is no level 3. In those cases, level 1 or 2 would be acceptable without justification. IMDRF Adverse Event codes Annex G (Component) will be published in 2020 as foreseen. The transition for its implementation in Europe is 12 months after its publication on the IMDRF website. | N | N | N | N | N
g | Description of remedial action/corrective action/preventive action / Field Safety Corrective Action | Describe any remedial, corrective and / or preventative action(s) taken as a result of this event and / or your investigation. | Y | N | N | Y | N
h | Time schedule for the implementation of the identified actions | Describe your time schedule for any corrective, preventative or other actions arising from this event and your investigation. | N | N | N | N | N
i | Final comments from the manufacturer on Cause investigation and conclusion | Enter your final comments on this incident for cause investigation and conclusion If you refer to an existing CAPA please provide the reference or identification number of that CAPA | N | N | N | N | N
4.3 | Similar Incidents
4.3.1 | Use of IMDRF terms and codes for identifying similar incidents
a | Identification of similar incidents using IMDRF Adverse Event Reporting terms and codes | Incidents occurring with the same device type/variant of a given manufacturer, and • having the same investigation finding (IMDRF investigation finding; Annex C) and the same medical device problem (IMDRF medical device problem; Annex A), if the investigation finding of the original incident is known, or • having the same medical device problem if the investigation finding is unknown; Notwithstanding the outcome for the patient, user or other person. When using the IMDRF Adverse Event Reporting terms and codes, it is preferable to identify and provide similar incident data based on the most relevant codes (this should be choice 1 in each coding section). Other - Where the IMDRF 'medical device problem' and 'investigation findings' codes are not used to describe similar incident data, please detail how the similar incidents data has been calculated and the rationale for not using the IMDRF codes specified. For example, for some implantable devices, it may be appropriate to provide similar incident data based on the IMDRF Health effect codes (Clinical signs, symptoms and conditions, Annex E; Health impact, Annex F). If using these health effect codes, please use the most relevant term from either Annex E or Annex F (and not a combination of both) for identifying similar incidents. This is to be used if the manufacturer decides to use the form for the transitional period (Dec 2018-Dec 2019) and has not yet implemented IMDRF codes. After Dec 2019 this section will not be used as similar incidents will be based on the IMDRF codes. | Y | N | N | Y | N
4.3.2 | Use of in-house terms/codes for identifying similar incidents (only for transition period)
a | If similar incident were not identified by IMDRF codes but by in-house codes, please indicate the code / combination of codes below. | Incidents occurring with the same device type/variant of a given manufacturer, and • having the same root cause evaluation and the same medical device problem, if the root cause evaluation of the original incident is known, or • having the same medical device problem if the root cause evaluation is unknown; Notwithstanding the outcome for the patient, user or other person. When using in-house codes and terms, it is preferable to identify and provide similar incident data based on the most relevant code and term. Other - Where the in-house 'medical device problem' and 'root cause evaluation' terms and codes are not used to describe similar incident data, please detail how the similar incidents data has been calculated and the rationale for not using these terms/codes. For example, for some implantable devices, it may be appropriate to provide similar incident data based on in-house patient problem terms and codes. This is to be used if the manufacturer decides to use the form for the transitional period (Dec 2018-Dec 2019) and has not yet implemented IMDRF codes. After Dec 2019 this section will not be used as similar incidents will be based on the IMDRF codes. | Y | N | N | Y | N
4.3.3 | Number of similar incidents and devices on the market
a | Indicate on which basis similar incidents were identified regarding the device or device variant | How have you defined and categorised similar incidents? Specify the particular products, batches, serial number range etc that you believe is the most appropriate for categorisation of relevant similar incidents. Pleae Note- single choice | Y | N | N | Y | N
b | Indicate to what criteria the number of devices on the market (also known as denominator data) is based on ( tick the most appropriate) | Indicate what the cumulative numbers are based on - for example number of devices sold or numbers of devices installed, number of tests performed; number of episodes of use for reusable devices (e.g. the number of episodes of use multiplied by the number of devices in service or sold). In cases of uncertainties which denominator data should be chosen, please discuss with the coordinating CA i.e. Manufacturers should work with the CA on the most appropriate denominator data for the device (or variant) concerned . In the future this information/guidance may be included in the DSVGs Please use the 'other- describe' if none of the above are suitable For a device that may receive events over its entire service life, the ideal denominator values are the number of remaining active devices within each given time period. However, since most countries do not report in-service and out of service data to manufacturers, provide the all-time cumulative sales from the date of the CE Mark approval to the end date of each given time period instead of within each given time period N.B- single choice | Y | N | N | Y | N
c | Enter the number of similar incidents and devices on the market for the indicated time periods | Enter, for the indicated time periods, the number of similar incidents and number of devices on market (i.e. denominator data should align with 4.3.3.b above). The default time period for similar incident data is yearly unless (a) a different time period has been specified by the European Medical Device Vigilance Experts Group or (b) the device has not been on the European market for more than three years. In the event of (b) please clearly indicate the time periods being used to provide similar incident data. If selecting yearly periods, the first selection will automatically set the other time periods | Y | N | N | Y | N
c | Start date | The default time period as described above-If a manufacturer believes there is good reason not to use the default time period the manufacturer should approach their CA requesting agreement to use the new reporting time periods with all Competent Authorities- in the future this information/guidance can be included in the DSVGs | Y | N | N | Y | N | Only Mandatory if using different dates from calendar years
c | End date | The default time period as described above-If a manufacturer believes there is good reason not to use the default time period the manufacturer should approach their CA requesting agreement to use the new reporting time periods with all Competent Authorities- in the future this information/guidance can be included in the DSVGs | Y | N | N | Y | N | Only Mandatory if using different dates from calendar years
c | Number of similar incidents that occurred within each time period: | Indicate the number of similar incidents. If none please add zero (type in "0") | Y | N | N | Y | N
c | Number of devices on market | Indicate the number of devices placed on market or the selected option for the denominator data in section 4.3.3.b. If none please add zero (type in "0") Please select the end of the previous months if data on numbers is not possible up to the date of the incident | Y | N | N | Y | N
c | In country where incident occurred | In the country where the incident occurred. If none please add zero (type in "0") | Y | N | N | Y | N
c | EEA + candidate countries + CH + TR | In EEA + CH + TR. If none please add zero (type in "0") This number will include the number directly above | Y | N | N | Y | N
c | World | In the world. If none please add zero (type in "0") This number will include the number directly above | Y | N | N | Y | N
d | Comments on how similar incidents and associated number of devices on the market were determined | Add any additional information here | N | N | N | N | N
| 5. General Comments | Please use this section for any further comments or information that has not already been provided in the earlier sections of this report form | N | N | N | N | N
| Coded Summary of report (Will be auto populated from previous selections) | This section will be auto populated from previous sections for a quick overview of the coded incident There is nothing to be filled out here by the manufacturer

## Risk Class and NB properties
| Combined initial & final | Initial | Follow Up | Final (Reportable incident) | Final (Non-reportable incident)
MDD/AIMDD IVDD
AIMD active implant | Y | N | N | Y | N
MDD class III | Y | N | N | Y | N
MDD class IIb | Y | N | N | Y | N
MDD class IIa | Y | N | N | Y | N
MDD class I | N | N | N | N | N
MDD class Is | Y (for part “sterile”) | N | N | Y (for part “sterile”) | N
MDD class Im | Y (for part “measuring function”) | N | N | Y (for part “measuring function”) | N
MDD class Ism | Y (for part "sterile" and part “measuring function”) | N | N | Y (for part "sterile" and part “measuring function”) | N
IVD Annex II List A | Y | N | N | Y | N
IVD Annex II List B | Y | N | N | Y | N
IVD devices for self-testing | Y | N | N | Y | N
IVD general | N | N | N | N | N
MDR/IVDR
MDR class III | Y | N | N | Y | N
MDR class IIb | Y | N | N | Y | N
MDR class IIa | Y | N | N | Y | N
MDR class I | N | N | N | N | N
MDR class I + “sterile conditions” | Y | N | N | Y | N
MDR class I + “measuring function” | Y | N | N | Y | N
MDR class I + “reusable surgical instruments” à yes | Y | N | N | Y | N
IVDR class D | Y | N | N | Y | N
IVDR class C | Y | N | N | Y | N
IVDR class B | Y | N | N | Y | N
IVDR class A | N | N | N | N | N
IVDR class A + "sterile conditions" | Y | N | N | Y | N

## Rules
Question | Default status | Description
1.3.1 a free text | disabled | Enabled if "Other, please specify" is selected
1.3.4 a-l | enabled | Autopopulated if the submitter of the report is a Manufacturer or Authorised representative
2.2 a free text | disabled | Enabled if "Other, please specify" is selected
2.3 q Year and Month | disabled | Enabled if "If software, date first made available" is selected
2.4 c MDR Type (Multiple choice) | disabled | Enabled if an MDR class is selected. Custom made is always enabled.
2.4 c IVDR Type (Multiple choice) | disabled | Enabled if an IVDR class is selected
3.2 c free text | disabled | Enabled if "Other" is selected
3.2 d free text | disabled | Enabled if "Other, please describe" is selected
3.2 e free text | disabled | Enabled if "Other" is selected
3.4 a free text | disabled | Enabled if "Other, please specify" is selected
4.2 d If 'No', rationale for no review required | disabled | Enabled if "No" ("Has the risk assessment been reviewed?") is selected
4.2 d If the risk assessment has been reviewed, is it still adequate? | disabled | Enabled if "Yes" ("Has the risk assessment been reviewed?") is selected
4.3.1 a Annex C (Choice 1) | disabled | Enabled if Annex A (of the same question) is selected
4.3.1 a free text | disabled | Enabled if "Other" is selected
4.3.2 a Annex C Code/Term | disabled | Enabled if "Code for most relevant medical device problem" (of the same question) is specified
4.32 a free text | disabled | Enabled if "Other" is selected
4.3.3 b free text | disabled | Enabled if "Other - describe" is selected

## Field Map
| Section | XML element
1. Administrative information
1.1 | Corresponding competent authority
a | Name of receiving national competent authority (NCA) | ncaName
b | EUDAMED Number of NCA | ncaEudamedNum
c | Reference number assigned by NCA | ncaReportNo
d | Reference number assigned by EUDAMED for this incident | refNumEudamed
1.2 | Date, type, and classification of incident report
a | Date of submission | reportDate
b | Date of incident | adverseEventDateFrom adverseEventDateTo
c | Manufacturer awareness date | mfrAwarenessDate
d | Type of Report | reportType
e | In case of initial and follow-up reports, please indicate the expected date of next report | reportNextDate
f | Classification of incident | eventClassification
1.3 | Submitter information
1.3.1 a | Submitter of report | statusReporter
b | Manufacturer's reference number for this incident | mfrRef
c | If this incident involves multiple devices from the same manufacturer, please list the respective reference numbers of the other MIR forms you have submitted | ncaRefMultiDev eudamedRefMultiDev mfrRefMultiDev
d | If this incident is covered under an FSCA, please provide the relevant numbers: | ncaRefFSCA eudamedRefFSCA mfrRefFSCA
e | Periodic Summary Report (PSR) ID | psrId
f | If the incident occurred within a PMCF/ PMPF investigation; please provide the Eudamed ID of that PMCF/PMPF investigation | pmcfpmpfId
1.3.2 | Manufacturer information
a | Manufacturer Organisation name | mfrName
b | Single Registration Number (SRN) | mfrSRN
c | Contact's first name | mfrContactPersonFirstName
d | Contact's last name | mfrContactPersonSecondName
e | E-mail | mfrEmailAddress
f | Phone | mfrPhone
g | Country | mfrCountry
h | Street | mfrOrgStreet
i | Street number | mfrOrgStreetNum
j | Address complement | mfrAddress
k | P.O Box | mfrOrgPOBox
l | City name | mfrCity
m | Postcode | mfrPostcode
1.3.3 | Authorized representative information
a | Authorised representative Organisation name | ARName
b | Single Registration Number (SRN) | ARSRN
c | Contact's first name | ARContactPersonFirstName
d | Contact's last name | ARContactPersonSecondName
e | E-mail | AREmailAddress
f | Phone | ARPhone
g | Country | ARCountry
h | Street | ARStreet
i | Street number | ARStreetNum
j | Address complement | ARAddress
k | P.O Box | ARPOBox
l | City name | ARCity
m | Postcode | ARPostCode
1.3.4 | Submitters details if not manufacturer or authorized representative
a | Registered commercial name of company | reporterOrgName
b | Contact's first name | reporterContactPersonFirstName
c | Contact's last name | reporterContactPersonSecondName
d | E-mail | reporterOrgEmailAddress
e | Phone | reporterOrgPhone
f | Country | reporterOrgCountry
g | Street | reporterOrgStreet
h | Street number | reporterOrgStreetNum
i | Address complement | reporterOrgAddress
j | P.O Box | reporterOrgPOBox
k | City name | reporterOrgCity
l | Postcode | reporterOrgPostcode
2. Medical Device Information
2.1 | Unique Device Identification (UDI)
a | UDI device identifier | udiDI
b | UDI production identifier | udiPI
c | Basic UDI-DI?eudamed DI | udiDIBasic
d | Unit of use UDI-DI | udiDIUnitUse
2.2 | Categorisation of device
a | Medical Device terminology used | nomenclatureSystem nomenclatureSystemOther
b | Medical device nomenclature code | nomenclatureCode
2.3 | Description of device and commercial information
a | Medical device name (Brand/Trade/Proprietary or Common name) | brandName
b | Nomenclature text/Description of the device and its intended use | deviceDescription
c | Model | modelNum
d | Catalogue/Reference number | catalogNum
e | Serial number(s) | serialNum
f | Lot/batch number(s) | batchNum
g | Software version | deviceSoftwareVer
h | Firmware version | deviceFirmwareVer
i | Device manufacturing date | deviceMfrDate
j | Device expiry date | deviceExpiryDate
k | Date when device was implanted | ImplantedDateFrom ImplantedDateTo
l | Date when device was explanted | ExplantedDateFrom ExplantedDateTo
m | If precise implant/explant dates are unknown, provide the duration of implantation | numYears numMonths numDays
n | Implant facility | implantFacilityName
o | Explant facility | explantFacilityName
p | Notified body (NB) ID number(s) (if applicable ) | nbCertNum, nbCertNum2
p | Notified body (NB) certificate number(s) of device (if applicable) | nbIdNum, nbIdNum2
q | Please indicate the date of one of the following: | deviceMarketDateType deviceMarketDate (year, month)
2.4 | Risk class of device when placed on market
a | This device has been placed on the market before the implementation of the MDD/AIMDD/IVDD | deviceClass
b | MDD/AIMDD/IVDD Class | deviceClassMDD deviceClassIVDD
b | MDR/IVDR Class | deviceClassMDR deviceClassIVDR deviceClassMDRType deviceClassIVDRType
2.5a | Market distribution of device (region/ country) (according to the best knowledge of the manufacturer) | distribution_all distribution otherCountries
2.6 | Use of accessories, associated devices or other devices
a | Relevant accessories used with the device being reported on, if applicable | deviceAccessories
b | Relevant associated devices used with the device being reported on, if applicable | deviceAssociated
3.Incident information derived from health care professional and/or user
3.1 | Nature of incident
a | Provide a comprehensive description of the incident, including (1) what went wrong with the device (if applicable) and (2) a description of the health effects (if applicable), i.e. clinical signs, symptoms, conditions as well as the overall health impact | eventDescription
3.2 | Medical device problem information
a | IMDRF Medical device problem codes (annex A)
b | Number of patients involved | numPatientsInvolved
c | What is the current location of the device | currentDeviceLocation currentDeviceLocationOtherText
d | Operator of device at the time of the incident | deviceOperatorAtEvent deviceOperatorAtEventOther
e | Usage of device | deviceUsage deviceUsageOther
f | Remedial actions taken by healthcare facility, patient or user subsequent to the incident | patientRemedialAction
3.3 | PATIENT INFORMATION
a | IMDRF health effect codes | imdrfClinicalCodeChoice1-6 imdrfHealthCodeChoice1-6 imdrfMissingCodeHealthEffect
b | Age of the patient at the time of incident, if applicable | numYearsPat numMonthsPat numDaysPat
c | Gender | Gender
d | Body Weight (Kg) | massKG
e | List any of the patient's prior health condition or medication that may be relevant to this incident | patientPriorMedication
3.4 | INITIAL REPORTER (can be healthcare professional of facility, patient, lay user)
a | Role of initial reporter | initialReporterRole initialReporterRoleOther
b | Name of the health care facility where incident occurred | healthcareFacilityName
c | Healthcare facility report Number | healthcareFacilityRepNum
d | Contact's first name | healthcareFacilityContactPersonFirstName
e | Contact's last name | healthcareFacilityContactPersonSecondName
f | Email | healthcareFacilityEmail
g | Phone | healthcareFacilityPhone
h | Country | healthcareFacilityCountry
i | Street | healthcareFacilityStreet
j | Street number | healthcareFacilityStreetNum
k | Address complement | healthcareFacilityAddress
l | P.O Box | healthcareFacilityPOBox
m | City name | healthcareFacilityCity
n | Postcode | healthcareFacilityPostcode
4. Manufacturer analysis
4.1 | Manufacturer’s preliminary comments
a | For initial and follow-up reports: preliminary results and conclusions of manufacturer’s investigation | manufacturersPrelimAnalysis
b | Initial actions (corrective and/or preventive) implemented by the manufacturer | manufacturersInitialCorrecAction
c | What further investigations do you intend in view of reaching final conclusions? | furtherInvestigations
4.2 | Cause investigation and conclusion
a | For Final (Reportable incident)- Description of the manufacturer’s evaluation concerning possible root causes/causative factors and conclusion | rootCauses
b | For Final (Non-reportable incident)- Fill out rationale for why this is considered not reportable | manufacturersWhyNotReportable
c | Is root cause confirmed | rootCauseConfirmed
d | Has the risks assessment been reviewed? | riskAssReviewed rationaleNoReview riskAssAdequate riskAssResults
e | IMDRF Cause investigation terms and codes | imdrfTypeInvestigationCodeChoice1-8 imdrfInvestigationFindingsCodeChoice1-6 imdrfInvestigationConclusionCodeChoice1-6 imdrfMissingCodeCauseInvestigation
f | IMDRF Component codes | imdrfComponentCodeChoice1-6 imdrfMissingComponentCodes
g | Description of remedial action/corrective action/preventive action / Field Safety Corrective Action | correctiveAction
h | Time schedule for the implementation of the identified actions | correctiveActionSchedule
i | Final comments from the manufacturer on Cause investigation and conclusion | manufacturersFinalComments
4.3 | Similar Incidents
4.3.1 | Use of IMDRF terms and codes for identifying similar incidents
a | Identification of similar incidents using IMDRF Adverse Event Reporting terms and codes | similarIncImdrfCodesAnnexA similarIncImdrfCodesAnnexC similarIncOtherReportingCodes similarIncOtherReportingCodesText
4.3.2 | Use of in-house terms/codes for identifying similar incidents (only for transition period)
a | If similar incident were not identified by IMDRF codes but by in-house codes, please indicate the code / combination of codes below. | similarIncMdCodesChoice1Code similarIncMdCodesChoice1Term similarIncRootCouseCodesChoice1Code similarIncRootCouseCodesChoice1Term similarIncOtherInHouseCodes similarIncOtherInHouseCodesText
4.3.3 | Number of similar incidents and devices on the market
a | Indicate on which basis similar incidents were identified regarding the device or device variant | similarVariant similarVariantDetails
b | Indicate to what criteria the number of devices on the market (also known as denominator data) is based on ( tick the most appropriate) | numberBasedOn numberBasedOnOther
c | Enter the number of similar incidents and devices on the market for the indicated time periods
c | Start date | startDateN - startDateN-3
c | End date | endDateN - endDateN-3
c | Number of similar incidents that occurred within each time period:
c | Number of devices on market
c | In country where incident occurred | countryIncidentIncN - countryIncidentIncN-3
c | EEA + candidate countries + CH + TR | eeaChTrIncN - eeaChTrIncN-3
c | World | worldIncN - worldIncN-3
d | Comments on how similar incidents and associated number of devices on the market were determined | AdditionalComments
5. General Comments
Coded Summary of report (Will be auto populated from previous selections)

## AnnexA Problem Code
Device (bold): For the purpose of this annex A, a device means a medical device including accessories and components.
Level 1 |  |  | Level 2 |  |  | Level 3
Term | Definition | Code | Term | Definition | Code | Term | Definition | Code
Patient Device Interaction Problem | Problem related to the interaction between the patient and the device. | A01 | Patient-Device Incompatibility | Problem associated with the interaction between the patient's physiology or anatomy and the device that affects the patient and/or the device. | A0101 | Biocompatibility | Problem associated with undesirable local or systemic effects due to exposure to medical device materials or leachates from those materials by a patient who has an implant or is receiving treatment with a device made from them. | A010101
[6 empty] | Device Appears to Trigger Rejection | The device appears to elicit undesired response in the patient to the presence of an implanted or invasive device, without inherent device failure, e.g. fibrous encapsulation, or inflammation of the tissue around the device, or extrusion of the device. | A010102
[6 empty] | Inadequacy of Device Shape and/or Size | The physical size and/or shape of the device was inadequate with regard to the patient's anatomy. | A010103
|  |  | Osseointegration Problem | Problem associated with interconnection between the bone tissue and the implanted device. | A0102 | Failure to Osseointegrate | Problem associated with the failure to see direct anchorage of an implant by the formation of bony tissue around the implant without the growth of fibrous tissue at the bone-implant interface. | A010201
[6 empty] | Loss of Osseointegration | Problem associated with weakened integration of the device at the bone-implant interface due to loss of fibrous and/or bony tissue and leading to compromised anchorage of the device. i.e. 'Loosening/Lysis.' " | A010202
|  |  | Loosening of Implant Not Related to Bone-Ingrowth | Problem associated with the loss of direct anchorage of an implanted device over time or due to an injury. | A0103
|  |  | Migration or Expulsion of Device | Problem with an implanted or invasive device moving within the body, or being completely expelled from the body. | A0104 | Expulsion | Problem with all or part of an implanted or invasive device being completely expelled from its intended location within the body. | A010401
[6 empty] | Migration | Problem with all or part of an implanted or invasive device moving from its intended location within the body. | A010402
Manufacturing, Packaging or Shipping Problem | Problem associated with any deviations from the documented specifications of the device that relate to nonconformity during manufacture to the design of an item or to specified manufacturing, packaging or shipping processes (out of box problem). | A02 | Product Quality Problem | Problem associated with an inherent device characteristic that is not satisfactory as specified or delivered. | A0201 | Dull, Blunt | Problem associated with a device not being as sharp as intended or expected. | A020101
[6 empty] | Nonstandard device | Problem associated with the device that does not meet the specifications or requirements for which it was manufactured (e.g. materials, parts, manufacturing process). | A020102
|  |  | Defective Component | Problem associated with a device component having flaws of dimensional deviations greater than acceptable for the intended use. | A0202
|  |  | Defective Device | Problem associated with having flaws or dimensional deviations greater than acceptable for the intended use of the device. | A0203
|  |  | Device Damaged Prior to Use | Problem associated with packaging or shipping damage prior to the use of the device. | A0204
|  |  | Packaging Problem | Problem associated with the materials used to construct the cover or outer wrapping of the device. | A0205 | Difficult to Open or Remove Packaging Material | Problem associated with difficulty for users to operate the device, specifically as it relates to the opening or removal of the outer wrapping. | A020501
[6 empty] | Incomplete or Missing Packaging | Problem associated with the nonconformance to the device specifications due to incomplete or missing packaging that may compromise the device operation as intended. | A020502
[6 empty] | Unsealed Device Packaging | Problem associated with the loss of packaging seal. | A020503
[6 empty] | Tear, Rip or Hole in Device Packaging | Problem associated with packaging damage (tear, rip or hole) prior to the use of the device. | A020504
|  |  | Device Misassembled During Manufacturing / Shipping | A device found incorrectly assembled when delivered to the user facility. | A0206 | Component Misassembled | A device found to have one or more components incorrectly assembled when delivered to the user facility. | A020601
[6 empty] | Component Missing | A device component(s) found to be missing when delivered to the user facility. | A020602
|  |  | Shipping Damage or Problem | Problem associated with shipping damage or problem prior to the use of the device. | A0207 | Delivered as Unsterile Product | Problem associated with a device being received in such a manner to indicate that its sterility has been compromised (e.g. sterile packaging breached, visible contaminate present) | A020701
Chemical Problem | Problem associated with any from the documented specifications of the device that relate to any chemical characterization, i.e., element, compound, or mixture. | A03 | Device Emits Odor | Problem associated with an unexpected or inappropriate smell released by the device. | A0301
|  |  | Device Ingredient or Reagent Problem | Problem associated with any deviations from the documented specifications of the device that relate to any ingredient or reagent characterization. | A0302 | Clumping in Device or Device Ingredient | Problem associated with the aggregation of particles into irregular masses. | A030201
[6 empty] | Coagulation in Device or device Ingredient | Problem associated with the undesired characterization of congealing, solidifying, thickening, curdling. | A030202
[6 empty] | Precipitate in Device or device Ingredient | Problem associated with the separation of solid particles from a liquid as the result of a chemical or physical change. | A030203
[6 empty] | Cross Reactivity | Problem associated with the degree to which an antibody or antigen participates in cross reactions. | A030204
[6 empty] | Particulates | Substances that consist of separate particles that are introduced by the device during use. | A030205
[6 empty] | High pH | pH higher than expected and / or anticipated. | A030206
[6 empty] | Low pH | pH lower than expected and / or anticipated. | A030207
|  |  | Improper Chemical Reaction | Problem associated with an unexpected or incomplete chemical reaction or effect. | A0303
Material Integrity Problem | Problem associated with any deviations from the documented specifications of the device that relate to the limited durability of all material used to construct device. | A04 | Break | Problem associated with undesired damage or breakage of those materials used in the device construction. | A0401 | Fracture | Problem associated with a partial or full-thickness crack in the device materials. | A040101
[6 empty] | Loss of or Failure to Bond | Problem associated with lack or loss of adherence between materials intended to be joined together by an adhesive. | A040102
[6 empty] | Material Fragmentation | Problem associated with small pieces of the device breaking off unexpectedly. | A040103
[6 empty] | Solder Joint Fracture | Problem associated with undesired damage or breakage in a solder joint of materials used in the device construction. | A040104
|  |  | Burst Container or Vessel | Problem associated with the pressure inside a vessel or container rising to such a degree that the container ruptures. | A0402
|  |  | Explosion | Problem associated with the violent bursting due to the sudden expansion of air, gas or fluid. | A0403
|  |  | Crack | Problem associated with an undesired partial separation and/or a visible opening along the length or width in the materials that are used in the device construction. | A0404
|  |  | Degraded | Problem associated with a undesired change in the chemical structure, physical properties, or appearance in the materials that are used in the device construction. | A0405 | Calcified | Problem associated with buildup of calcium salts on the device. | A040501
[6 empty] | Corroded | Problem associated with the chemical or electrochemical reaction between materials, usually a metal and its environment that produces a deterioration of the metal and its properties. | A040502
[6 empty] | Material Erosion | Problem associated with a progressive loss of a material from a solid surface. | A040503
[6 empty] | Pitted | Problem associated with the corrosion of a material's surface, confined to a point or small area that takes the form of cavities. | A040504
[6 empty] | Flaked | Problem associated with the detachment of small pieces of the coating film of a material. | A040505
[6 empty] | Peeled / Delaminated | Peeling or delamination of composite materials, including coatings, that occurs when layers are separated as a result of stress or impact and resulting in loss of mechanical toughness. | A040506
[6 empty] | Naturally Worn | Problem associated with material damage to a surface, usually involving progressive loss or displacement of material, due to relative motion between that surface and a contacting substance or substances. | A040507
[6 empty] | Unraveled Material | Problem due to the undesired unravelling of material (e.g. disentangled, unwound etc.). | A040508
|  |  | Material Deformation | Problem associated with an undesired material change in shape or property caused by external forces. | A0406 | Deformation Due to Compressive Stress | Problem associated with an undesired bulge, bend, bow, kink, or wavy condition observed in the device material resulting from compressive stresses. | A040601
[6 empty] | Dent in Material | Problem associated with a undesired change in shape, characterized by the presence of a slight hollow (dent) in the device surface. | A040602
[6 empty] | Failure to Fold | Problem associated with an undesired material change in physical property, characterized by failure to fold. | A040603
[6 empty] | Failure to Unfold or Unwrap | Problem associated with the comprising materials' deformation in that device fails to open its wrapping or open/extend in a certain manner i.e. balloon or lens. | A040604
[6 empty] | Material Frayed | Problem associated with the comprising materials having damaged edges. | A040605
[6 empty] | Material Invagination | Problem associated with an undesired material change in shape, characterized by the infolding of one part within another part of a structure. | A040606
[6 empty] | Material Too Rigid or Stiff | Problem associated with an undesired material change in physical property, characterized by rigidity (it resists deformation in response to an applied force). | A040607
[6 empty] | Material Too Soft / Flexible | Problem associated with any device material that results in the material's inability to maintain the desired shape or support function. | A040608
[6 empty] | Material Twisted / Bent | Problem associated with deformations that lead to twisting or bending of the device. | A040609
[6 empty] | Melted | Problem associated with a solid device being transformed into a molten or liquid state. | A040610
[6 empty] | Stretched | Problem associated with an increase or elongation in a materials' dimension. | A040611
|  |  | Material Discolored | Problem associated with an undesired streak, pattern and/or a noticeable change in color from the rest of the materials used in the device construction. | A0407
|  |  | Material Disintegration | Problem associated with material breaking into small particles. | A0408
|  |  | Material Opacification | C62895; FDA 1426 - Problem associated with an undesirable opaqueness or cloudiness. | A0409
|  |  | Material Perforation | Material constituting device is perforated possibly compromising the device's intended purpose. | A0410 | Material Puncture / Hole | Device material(s) punctured leading to undesired holes/openings. | A041001
|  |  | Material Protrusion / Extrusion | Problem associated with undesired physical appearance of device material, specifically when material extends beyond or above device surface. | A0411
|  |  | Material Rupture | Problem associated with perforations that lead to bursting of the device. | A0412
|  |  | Material Separation | Problem associated with an undesired disassociation or breaking apart of the device. | A0413
|  |  | Material Split, Cut or Torn | Problem associated with materials consisting the device are split, cut or torn due to external forces (e.g. wrenching or laceration) or internal forces (e.g. exceeding the tensile stress limits belonging to the materials used in the device construction). | A0414
|  |  | Scratched Material | Problem associated with an undesirable shallow cut or narrow groove in the surface of the device materials. | A0415
Mechanical Problem | Problems associated with mechanical actions or defects, including moving parts or subassemblies, etc. | A05 | Detachment of Device or device Component | Problem associated with the separation of the device from its physical construct, integrity, or chassis. | A0501
|  |  | Device Damaged by Another Device | Problem associated with one device causing harm to another device. | A0502
|  |  | Ejection Problem | Problems associated with the inability of or unexpected removal or separation of the device from its physical location. | A0503 | Failure to Eject | Problem associated with the inability to remove or discharge the device from the location of use. | A050301
[6 empty] | Unintended Ejection | Problem associated with unexpected discharge of the device from expected location includes but not limited to the device such as clip appliers, film cartridge, staples. | A050302
|  |  | Leak / Splash | Problem associated with the escape of a liquid or gas from the vessel or container in which it is housed. | A0504 | Fluid Leak | Escape (Release, Discharge) of fluid through an unintended location - often accompanied by a loss of pressure and/or output. | A050401
[6 empty] | Gas Leak | Problem associated with the unintended escape of a gas from the container in which it is housed. | A050402
[6 empty] | Gel Leak | Escape (Release, Discharge) of gel through an unintended location - as in leakage of ultrasound gel. Escape or release of gel from containment structures - as in gel filled implant leak. | A050403
[6 empty] | Radiation Leak | Escape of radiation (energy in the form of waves or subatomic particles, especially those that cause ionization) through containment structures, leading to unintended exposure. | A050404
[6 empty] | Perivalvular Leak | Problem associated with the escape of blood around a heart valve, particularly around its leaflets. | A050405
|  |  | Firing Problem | Problem associated with the device not discharging as intended. | A0505 | Failure to Fire | Problem associated with failure of the device to discharge its load (e.g. surgical stapler failed to partially or completely deploy its staples). | A050501
[6 empty] | Misfire | Problem associated with a therapy or algorithm not being delivered or executed at the expected time. | A050502
|  |  | Mechanical Jam | The motion of the device is prevented or restricted. | A0506
|  |  | Mechanics Altered | Problem associated with a device mechanical functioning of machinery, moving parts or tools of device being changed or modified. | A0507 | Failure to Align | Problem associated with a circuit, equipment, or system whereby its functions fail to be properly synchronized or its relative positions properly oriented. | A050701
[6 empty] | Failure to Cut | Inability of the device to make an incision, pierce or open as intended. | A050702
[6 empty] | Failure to Cycle | Problem associated with the device failing to complete a series of processes or events. | A050703
[6 empty] | Failure to Form Staple | Problem associated with the device failing to connect tissue with a stapling device due to the staples not forming correctly. | A050704
|  |  | Noise, Audible | Problem associated with any unintended sound which emanates from the device (for example, squeaking from two parts rubbing together or buzzing sounds from electrical components). | A0508
|  |  | Physical Resistance / Sticking | Problem associated with the lack of movement in the device due parts sticking or seizing. | A0509
|  |  | Retraction Problem | Problem associated with drawing back the device to an intended location. | A0510
|  |  | Structural Problem | Problem associated with the basic physical construction or physical make up of the device. | A0511 | Collapse | Problem associated with the buckling or crushing of material from external forces. | A051101
[6 empty] | Sharp Edges | The device has undesirable sharp edges which can cause harm or damage. | A051102
[6 empty] | Difficult to Fold or Unfold | Problem associated with the use of the device in terms of user experiencing difficulty to close or to spread out/extend length of the device, even if the operation is being performed according to labeled instructions for use. | A051103
[6 empty] | Difficult to Open or Close | Problem associated with the use of the device in terms of user experiencing difficulty opening and closing the device, even if the operation is being performed according to labeled instructions for use. | A051104
[6 empty] | Incomplete Coaptation | Problem associated with the heart valve leaflet not closing properly. | A051105
|  |  | Unintended Movement | Problem associated with an undesired movement of the device, which may be related to the device- malfunction, misdiagnosis, or mistreatment. | A0512 | Device Dislodged or Dislocated | Problems associated with the device not remaining in an expected location. | A051201
[6 empty] | Device Tipped Over | Problem associated with the inability of the device to stay in an upright position. | A051202
[6 empty] | Device Fell | Problem associated with the device or a component unexpectedly being dropped or moving down from an intended place. | A051203
[6 empty] | Device Slipped | Problem associated with the device moving or sliding from the intended position. | A051204
[6 empty] | Unintended Collision | Problem associated with the device impacting with another object. | A051205
[6 empty] | Unintended System Motion | Problem associated with any motion of the system or components that was not initiated by the user. | A051206
[6 empty] | Unstable | Problem associated with the mechanical stability of the device. | A051207
[6 empty] | Vibration | Problem associated with the undesirable mechanical oscillation. | A051208
Optical Problem | Problem associated with transmission of visible light affecting the quality of the image transmitted or otherwise affecting the intended application of the visible light path. | A06 | Misfocusing | The problem relates to the poor focusing of the object or the focus is on the wrong object or in the wrong area. | A0601
|  |  | Optical Decentration | Problem associated with being off-center of optical lenses. | A0602
|  |  | Optical Discoloration | Problem associated with an undesired change of color. | A0603
|  |  | Optical Distortion | Problem associated with an optical defect in an image-forming system whereby the image is not the shape of an ideal image of the object. | A0604
|  |  | Optical Obstruction | Problem associated with the blocking of optical devices, e.g. visual pathways. | A0605
Electrical /Electronic Property Problem | Problem associated with a failure of the electrical circuitry of the device. | A07 | Capturing Problem | Problem associated with the inability of the device to achieve successful depolarization and contraction of a cardiac chamber caused by a pacemaker output pulse. | A0701 | Failure to Capture | Problem associated with the failure to achieve effective and consistent depolarization of the heart resulting from the electrical stimulus of the pacemaker. | A070101
[6 empty] | High Capture Threshold | Problem with the amount of output energy needed to cause cardiac depolarization being higher than expected/desired. | A070102
[6 empty] | Intermittent Capture | Problem associated with the ineffective and inconsistent depolarization of the heart. | A070103
[6 empty] | Unstable Capture Threshold | Problem with the amount of output energy needed to cause cardiac depolarization being unstable. | A070104
|  |  | Continuous Firing | Problem associated with the excessive production of electrical impulses over a period. | A0702
|  |  | Arcing | Problem associated with electrical current flowing through a gap between two conductive surfaces, typically resulting in a visible flash of light. | A0703 | Arcing at Paddles | Problem associated with electrical current flowing through a gap between paddles (conductive surfaces), typically resulting in a visible flash of light. | A070301
[6 empty] | Arcing of Electrodes | Problem associated with electrical current flowing through a gap between electrodes (conductive surfaces), typically resulting in a visible flash of light. | A070302
|  |  | Sparking | Problem associated with a flash of light related to an electrical discharge into a normally non conductive medium, such as air. Not associated with a discharge between two conductive surfaces. | A0704
|  |  | Battery Problem | Problem associated with the internal power of the device (e.g. battery, transformer, fuel cell or other power sources). | A0705 | Battery Problem: High Impedance | Problem related to increased battery internal impedance. | A070501
[6 empty] | Battery Problem: Low Impedance | Problem related to decreased battery internal impedance. | A070502
[6 empty] | Failure to Run on Battery | Problem associated with the device failing to operate when not connected to a fixed power source. | A070503
[6 empty] | Premature Discharge of Battery | Battery discharging earlier than expected. | A070504
|  |  | Charging Problem | Problem associated with the inability of the device to successfully charge an electrical source. | A0706 | Aborted Charge | Problem associated with the premature ending of the charging process (e.g. of a battery or other charge storage device). | A070601
[6 empty] | Delayed Charge Time | Problem associated with an unexpected amount of time required to charge the device (e.g. a delay in starting charging or a longer than expected charge time). | A070602
[6 empty] | Failure to Charge | Problem associated with inability to initiate the appropriate charging process (e.g. of a battery or other charge storage device) | A070603
|  |  | Failure to Discharge | Problem associated with the failure of a battery or other charge storage device to appropriately discharge as intended. Does not apply to defibrillation. | A0707
|  |  | Power Problem | Problem associated with the energy to operate the device. | A0708 | Complete Loss of power | Problem associated with the lack of power to run the device. | A070801
[6 empty] | Intermittent loss of power | Problem associated with an intermittent disruption to the power to run the device. | A070802
[6 empty] | Failure to Power Up | Problem associated with the inability of the device to turn on related to energy delivered to the device. | A070803
[6 empty] | Unintended power up | Problem associated with the device turning on when not intended. | A070804
|  |  | Device Sensing Problem | Problem associated with the device feature that are designed to respond to a physical stimulus (temperature, illumination, motion, cardiac rhythms) and that do not transmit a resulting signal for interpretation or measurement. | A0709 | Decreased Sensitivity | Problem with the device being less sensitive to an input than intended or expected. | A070901
[6 empty] | Increased Sensitivity | Problem with the device being more sensitive to an input than intended or expected. | A070902
[6 empty] | Failure to Analyze Signal | Problem with the device not analyzing a signal. | A070903
[6 empty] | Failure to Select Signal | Problem associated with the failure of the device to select the appropriate input signal. | A070904
[6 empty] | High Sensing Threshold | Problem associated with the amount of input required by the device to detect a signal being higher than expected/desired. | A070905
[6 empty] | Low Sensing Threshold | Problem associated with the amount of an input required by the device to detect a signal being lower than expected/desired. | A070906
[6 empty] | Loss of Threshold | Problem associated with the loss of the minimum amount of energy, voltage, or current needed to consistently stimulate the heart muscle. | A070907
[6 empty] | Failure to Sense | Problem associated with the failure of the device designed to respond to a physical stimulus (as temperature, illumination, motion) to transmit a resulting signal for interpretation or measurement. | A070908
[6 empty] | Over-Sensing | Problem related to failure of the device to properly filter cardiac signals resulting in inappropriate device response. | A070909
[6 empty] | Under-Sensing | Problem related to failure of the device to properly detect intrinsic cardiac activity and respond appropriately. | A070910
[6 empty] | Sensing Intermittently | Problem with the device receiving an incoming signal on an intermittent basis when expected to be continuous. | A070911
[6 empty] | Incorrect Interpretation of signal | Problem with the device inappropriately analyzing a signal. | A070912
|  |  | Failure to Conduct | Problem associated with the inability of the device to allow a current of electricity to pass or to conduct electricity continuously along an electrical path. | A0710
|  |  | Interrogation Problem | Problems associated with the device's ability to respond to signals from a system designed to interrogate its status. | A0711 | Difficult to Interrogate | Problem associated with difficulty of a transponder system to trigger a response. | A071101
[6 empty] | Failure to Interrogate | Problem associated with the device failure to appropriately respond to signals from a system designed to interrogate its status. | A071102
|  |  | Pacing Problem | Problem associated with the inability of the device to generate a therapeutic simulated heart beat via electrical impulses. | A0712 | Failure to Convert Rhythm | Failure of the device therapy or set of therapies to terminate the harmful cardiac rhythm that the therapy is meant to terminate. | A071201
[6 empty] | Inaccurate Synchronization | Problem associated with an error due to imperfect timing of two operations, e.g. signal transmission time. | A071202
[6 empty] | Inappropriate waveform | Failure of the device to generate a correctly-shaped pacing output, e.g., a waveform that is too wide. | A071203
[6 empty] | No Pacing | Problem associated with the device ceasing to deliver paces. | A071204
[6 empty] | Pacemaker Found in Back-Up Mode | A device with a pacing function found in back-up Mode. This may be an appropriate fail-safe action (e.g. end of battery life), or be caused by device-malfunction or due to operator error. | A071205
[6 empty] | Pacing Asynchronously | Problem associated with a pacing transmission process such that between any two significant instants in the same group, there is always an integral number of unit intervals. Between two significant instants located in different groups, there are not always an integral number of unit intervals. | A071206
[6 empty] | Pacing Inadequately | Pacing voltage or pulse width is less than desired. | A071207
[6 empty] | Pacing Intermittently | Problem associated with the failure of pacing device for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such as failure is often recurrent. | A071208
[6 empty] | Pocket Stimulation | Problem associated with a pocket of skin in which the pulse generator is housed. | A071209
|  |  | Defibrillation Problem | Problem associated with the inability of the device to provide an appropriate or successful electrical shock. | A0713 | Failure to Deliver Shock | Problem associated with the failure of the device to deliver electrical energy intended to change an electrical rhythm. | A071301
[6 empty] | Inappropriate Shock | Problem associated with the inappropriate delivery of an electrical energy. | A071302
[6 empty] | Intermittent Shock | Problem associated with the failure to deliver shock for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such as failure is often recurrent. | A071303
|  |  | Unintended Electrical Shock | The device delivers unintended electrical shock. Unintended defibrillation shock should be coded as "Inappropriate Shock" code number A071302. | A0714
|  |  | Grounding Malfunction | Problem associated with the inability to connect conductors of an electronic system for the purpose of controlling or impeding ground currents and voltages. | A0715
|  |  | Electrical Overstress | Problem associated with an electrical activity that exceeded the specified threshold limit of the internal integrated circuitry. | A0716
|  |  | Electro-Static Discharge | Problem associated with the discharge of electricity between two bodies previously electrically charged. | A0717
|  |  | Failure to Shut Off | Problem associated with the device not powering off when a shut down was requested. | A0718 | Device Remains Activated | Problem associated with the device continuing to be in an active state after deactivation was requested. | A071801
|  |  | Unexpected Shutdown | Problem associated with the device unexpectedly powering down. | A0719
|  |  | Electromagnetic Compatibility Problem | Problem associated with the ability of a system to function in its electromagnetic environment without introducing intolerable disturbances to anything in its environment. | A0720 | Electromagnetic Interference | Problem associated with a measure of electromagnetic radiation from equipment. | A072001
[6 empty] | Radiofrequency Interference (RFI) | Problem associated with the degradation of the reception of a wanted signal caused by RF disturbance. | A072002
|  |  | Circuit Failure | Problem associated with a failure of the internal network paths or electrical circuitry (i.e. electrical components, circuit boards, wiring) | A0721 | Capacitative Coupling | Problem associated with the transfer of energy within an electrical network by means of the capacitance between circuit nodes. It occurs when energy is coupled from one circuit to another through an electric field. | A072101
[6 empty] | Electrical Shorting | Problem associated with an electric current travelling along an accidental path (unintended path) in a circuit. | A072102
[6 empty] | Intermittent Continuity | Problem associated with intermittent faults in electrical/electronic interconnections. | A072103
|  |  | Impedance Problem | Problem associated with electrical impedance levels between device and patient connections. | A0722 | High impedance | Problem associated with higher than intended electrical impedance levels between device and patient connections. | A072201
[6 empty] | Low impedance | Problem associated with lower than intended electrical impedance levels between device and patient connections. | A072202
Calibration Problem | Problem associated with the operation of the device, related to its accuracy, and associated with the calibration of the device. | A08 | Failure to Calibrate | Problem associated with the failure of the device to perform a self-calibration procedure or process designed to assure the accuracy and proper performance of the device. | A0801
|  |  | Failure to Recalibrate | Problem associated with the failure of the device which is unable to regain a standard level of accuracy when performing a calibration procedure or process designed to assure the accuracy and proper performance of the device. | A0802
|  |  | Imprecision | Problem associated with the device providing imprecise measurements when compared to a reference standard. | A0803
|  |  | Overcorrection | Problem associated with an adjustment that surpasses a set of criteria. | A0804
Output Problem | Problem associated with any deviation from the documented specifications of the device that relate to the end result, data, or test results provided by the device. | A09 | Audible Prompt / Feedback | Problem with any deviation from the documented specifications of the device that relate to audible feedback. e.g. voice prompts or beeps, but not safety related alarms which are covered under "Protective Measures Problem". | A0901 | Inappropriate Audible Prompt / Feedback | Problem with audible messages which do not guide a device user to the correct action. | A090101
[6 empty] | Inaudible or Unclear Audible Prompt / Feedback | Problem associated with audible prompts which cannot be heard clearly. | A090102
[6 empty] | No Audible Prompt / Feedback | Problem associated with the device ceasing to provide audible prompts. | A090103
|  |  | Display or Visual Feedback Problem | Problem with any deviation from the documented specifications of the device that relate to visual feedback. e.g. the display of information, images on a screen, or output from the device. | A0902 | Device Displays Incorrect Message | Problem associated with providing incorrect display information. | A090201
[6 empty] | Display Difficult to Read | Problem associated with legibility of the display, compromising for instance the reading/interpretation of patient parameters or test results. Legibility problems can be due to color, size of font, display screen contrast or other factors. | A090202
[6 empty] | Erratic or Intermittent Display | A device does not consistently display the same message, result, reading, or image. e.g. the display might flicker, switch between readings or messages, or go completely blank for brief periods of time. | A090203
[6 empty] | Image Display Error / Artifact | Problem with image display leading to corrupted images or readouts/measurement indications. | A090204
[6 empty] | Image Orientation Incorrect | Problem associated with an incorrect image orientation on the device display. | A090205
[6 empty] | No Display / Image | Problem associated with the absence of display or image. | A090206
[6 empty] | No Visual Prompts / Feedback | Problem associated with the device ceasing to provide visual feedback. | A090207
[6 empty] | Poor Quality Image | Inadequate quality of an image or any visual representation displayed by the device, or output from the device. | A090208
[6 empty] | Visual Prompts will not Clear | Problem with visual messages which continue to be displayed on / by the device after the appropriate action has been taken. | A090209
|  |  | Tactile Prompts / Feedback | Problem with any deviation from the documented specifications of the device that relate to tactile feedback. e.g device vibrational prompt. | A0903 | Inappropriate Tactile Prompt / Feedback | Problem with tactile feedback which does not guide a device user to the correct action. | A090301
[6 empty] | No Tactile Prompts / Feedback | Problem associated with the device ceasing to provide tactile feedback. | A090302
|  |  | Energy Output Problem | Problem with the device's intended output of energy. | A0904 | Energy Spectrum Incorrect | Problem associated with the energy output from the device not being in the expected part of the spectrum. | A090401
[6 empty] | Failure to Deliver Energy | Problem associated with the failure of the device to deliver any energy. | A090402
[6 empty] | Intermittent Energy Output | Problem associated with the energy output from the device being inconsistent over time. | A090403
[6 empty] | Output above Specifications | Device output is exceeding the documented specifications of the device. | A090404
[6 empty] | Output below Specifications | Device output is below the documented specifications of the device. | A090405
[6 empty] | Therapeutic or Diagnostic Output Failure | Problem associated with the failure of the device to deliver the output required for treatment or identification of a disease. | A090406
[6 empty] | Therapy Delivered to Incorrect Body Area | Problem associated with the device causing unintended therapeutic action to an area of the body other than the intended area. | A090407
|  |  | Radiation Output Problem | Problem with the device's intended output of radiation. | A0905 | Radiation Output Failure | Problem associated with the absence of radiation output from radiological or diagnostic devices. | A090501
[6 empty] | Radiation Overexposure | Problem associated with excessive radiation emitted from radiological or diagnostic devices. | A090502
[6 empty] | Radiation Underexposure | Problem associated with too little radiation emitted from radiological or diagnostic devices. | A090503
[6 empty] | Unexpected/Unintended Radiation Output | Device-emitted radiation when it was not supposed to. This applies to devices which are intended to emit radiation, and the radiation being emitted from the correct part of the device, but at an incorrect time. Use "radiation leak" if the device emits radiation which should never have been emitted, or from a location from which it should never be emitted. | A090504
|  |  | Gas output problem | Problem associated with gas output. | A0906
|  |  | No Device Output | Problem associated with no measurement outcome, value or data obtained from the device. | A0907
|  |  | Incorrect, Inadequate or Imprecise Result or Readings | Problem associated with a nonconforming end result, data, or test results provided by the device to its performance specifications. | A0908 | Signal Artifact | Problem associated with impurities or interference in a signal (e.g. ECG artifact). | A090801
[6 empty] | Failure to Obtain Sample | The device does not collect or transfer the sample. | A090802
[6 empty] | False Negative Result | Problem associated with the device incorrectly reporting that something has not been detected and may mislead the operator into not taking certain actions when action should be taken. | A090803
[6 empty] | False Positive Result | Problem associated with the device incorrectly reporting that something has been detected and may mislead the operator to take certain actions. | A090804
[6 empty] | Incorrect Measurement | Measurement obtained from or provided by the device is obviously incorrect. | A090805
[6 empty] | Non Reproducible Results | Device results cannot be reliably reproduced. | A090806
[6 empty] | High Readings | Reading provided by the device is too high or higher than expected. | A090807
[6 empty] | Low Readings | Reading provided by the device is too low or lower than expected. | A090808
[6 empty] | High Test Results | Test results provided by the device are too high or higher than expected. | A090809
[6 empty] | Low Test Results | Test results provided by the device are too low or lower than expected. | A090810
[6 empty] | Unable to Obtain Readings | The device does not provide or display a valid reading. | A090811
[6 empty] | Missing Test Results | Problem associated with the results of a test or measurement not appearing. | A090812
|  |  | Unexpected Therapeutic Results | Problem associated with the use of the device for therapeutic purposes. | A0909
Temperature Problem | Problem associated with the device producing unintended temperatures. | A10 | Excessive Cooling | Problem associated with the device producing temperatures that are lower than specified. | A1001
|  |  | Excessive Heating | Problem associated with the device which have a warming or heating function, producing excessive heat. | A1002
|  |  | Insufficient Cooling | Problem associated with the device insufficiently cooled in device active (working) or/and non-active (nonworking) state. | A1003
|  |  | Insufficient Heating | Problem associated with the device or its components producing temperatures that are not as high as what is specified. | A1004
|  |  | Overheating of Device | Problem associated with the device producing high temperatures, such that its operation is compromised or harm is caused (e.g. overheating that produces melting of components or automatic shutdown). | A1005
|  |  | Thermal Decomposition of Device | Problems associated with a discoloration or destruction as a result of thermal decomposition of the device. | A1006
|  |  | Fire | Problem associated with the combustion of the device with a steady flame. | A1007
|  |  | Flare or Flash | Problem associated with device-related burn with an unsteady flame. | A1008
|  |  | Smoking | Problem associated with a cloud of vapor or gas generated from the device, generally associated after a fire or a burn. | A1009
Computer Software Problem | Problem associated with written programs, codes, and/or software system that affects device performance or communication with another device. | A11 | Application Network Problem | Problem associated with the deviations from documented system specifications that affects overall system performance and/or the performance of an individual device connected to that system. | A1101
|  |  | Application Program Problem | Problem associated with the requirement for software to fulfill its function within an intended use or application. | A1102 | Application Program Freezes, Becomes Nonfunctional | Problem associated with freezing and becoming nonfunctional of an application program. | A110201
[6 empty] | Application Program Problem: Dose Calculation Error | Problem associated with the written program code or application software used by the device to calculate specific measurements or quantities managed by the device. | A110202
[6 empty] | Application Program Problem: Medication Error | Event in which the device software results in errors of medication preparation or administration. | A110203
[6 empty] | Application Program Problem: Parameter Calculation Error | Problem associated with the written program code or application software used by the device to calculate parameters other than those related to dose or power. | A110204
[6 empty] | Application Program Problem: Power Calculation Error | Problem associated with the written program code or application software used by the device for calculations related to device power. | A110205
[6 empty] | Application Program Version or Upgrade Problem | Problem associated with installing updates to a software system that affects the device performance or communication with another device. | A110206
[6 empty] | Problem with Software Installation | Problem associated with installing the device software in a manner that allows full functioning of the device. Source of installation could be manufacturer or user. | A110207
[6 empty] | Unintended Application Program Shut Down | Problem associated with an unintended shut down by malfunction of the application program. | A110208
|  |  | Program or Algorithm Execution Problem | Problem associated with execution problems relating to program or algorithm. | A1103 | Delayed Program or Algorithm Execution | Problem associated with delayed execution relating to program or algorithm. | A110301
[6 empty] | Intermittent Program or Algorithm Execution | Problem associated with intermittent execution relating to program or algorithm. | A110302
[6 empty] | Program or Algorithm Execution Failure | Problem associated with the failure of a program or algorithm to execute. Sudden/unexpected interruption to a program's execution. | A110303
|  |  | Computer Operating System Problem | Problem associated with software, firmware, and/or hardware elements that control the execution of computer programs and provides such services as computer resource allocation, job control, input/output control, and file management in a computer system. | A1104 | Operating System Becomes Nonfunctional | Problem associated with malfunction of the computer operating system as opposed to an application software problem. | A110401
[6 empty] | Operating System Version or Upgrade Problem | Problem associated with replacing an older operating system to an up-to-date operating system. | A110402
|  |  | Computer System Security Problem | Problem associated with unauthorized access to or modification of a software system resulting in a loss of confidentiality, integrity, or availability of written program code, application software, or data or entire device. | A1105 | Application Security Problem | Problem associated with the acquisition of computer programming codes that can replicate and spread from one computer system to another thereby leading to damaged software, hardware and data. | A110501
[6 empty] | Unauthorized Access to Computer System | Problem associated with an access that was not permitted to the computer system that may lead to modification of program, corruption of data, or and break in network security. This concept is closely associated with computer integrity which is the degree to which a system or component prevents unauthorized access to, or modification of, computer programs or data. | A110502
|  |  | Data Back-Up Problem | Problems relating to a system, component, file, procedure, or person available to replace or help restore a primary item in the event of a failure or externally caused disaster. | A1106 | Failure to Back-Up | Problem associated with the inability to backup or to retrieve a backed up version (corrupted file) of device data or system files. | A110601
[6 empty] | Failure to Convert to Back-Up | Problem associated with a failure to transition from a primary system, component, file, procedure to a backup in response to a failure in the primary item. | A110602
|  |  | Data Problem | Event in which data (charting, orders, results) is not correctly stored, transferred, updated, or displayed. | A1107 | Loss of Data | Event in which data is unintentionally permanently or temporarily lost, deleted, corrupted, or overwritten. | A110701
[6 empty] | Patient Data Problem | Event in which data is accessed by the healthcare provider and either the wrong patient or the wrong data is retrieved despite correct inquiry procedures. | A110702
|  |  | Date/Time-Related Software Problem | Problem associated with programming of calendar dates and/or time as a factor in the operation of the device. | A1108
Connection Problem | Problem associated with linking of the device and/or the functional units set up to provide means for a transfer of liquid, gas, electricity or data. | A12 | Blocked Connection | Problem associated with linking of the device whereby their functional units set up to provide means for a transfer of fluid, gas, or data is prevented or impeded. | A1201
|  |  | Decoupling | Problem associated with the device being unassociated in such a way that fluid, gas, power or signal information may not be transferred from one to another. | A1202
|  |  | Disconnection | Problem associated with the linking of the device having a sufficient open space to prevent gas, liquid or electrical current flow between connectors. | A1203
|  |  | Failure to Disconnect | Problem associated with the linking of the device whereby termination of the transfer of liquid, gas, electricity, or information cannot be accomplished, or linking components do not come apart, or disconnect, when expected. | A1204
|  |  | Loose or Intermittent Connection | Problem associated with the connection of the device being loose or intermittent. | A1205
|  |  | Misconnection | Problem associated with the connection of the device being improper or not in accordance with device specification, requirements or intended uses. | A1206
|  |  | Incomplete or Inadequate Connection | Problem associated with a partial linking of the device whereby device may appear to be connected however only a partial, intermittent or no transfer of liquid, gas, electricity, or information can be accomplished. | A1207
|  |  | Fitting Problem | Problem associated with the connection of the device whereby channels, switching systems, and other functional units set up to provide means for a transfer of liquid, gas, electricity, or information do not match or fit. | A1208
Communication or Transmission Problem | Problem associated with the device sending or receiving signals or data. This includes transmission among internal components of the device to which the device is intended to communicate. | A13 | Failure to Read Input Signal | Problem associated with a failure of the device to read a signal for interpretation or measurement. | A1301
|  |  | Failure to Transmit Record | Problem associated with a failure of the device to transmit a record for interpretation or measurement. | A1302
|  |  | Intermittent Communication Failure | Inconsistent or lack of intended communication of data among internal components or with other external devices. | A1303
|  |  | Telemetry Discrepancy | Problem associated with variability of the transmission of telemetry signals. | A1304
|  |  | Wireless Communication Problem | Problems with the RF wireless technology characteristics and performance (e.g., frequency, output power, range, reception), wireless quality of service, wireless coexistence, security of wireless signals and data, and electromagnetic compatibility. | A1305
Infusion or Flow Problem | Problem associated with the device failing to deliver liquids or gases as intended (e.g. delivering drugs at incorrect rate, Problems with drawing fluid from a system.) | A14 | Deflation Problem | Problem associated with the inability of the device to release its contents. | A1401
|  |  | Excess Flow or Over-Infusion | Problem associated with a delivery overdose of therapeutic agents, such as drugs or fluids being delivered into a device or a patient. | A1402
|  |  | Filling Problem | Problem associated with the method or amount of time associated with the delivery of a fluid. Time to delivery or amount of delivered entity may be affected. | A1403 | Inability to Auto-Fill | Complete failure to fill as part of an automated process. For insufficient filling use "Short Fill". For excessive filling use "Overfill". For inconsistent filling use "Volume Accuracy Problem". | A140301
[6 empty] | Overfill | Excessive filling of a device. For complete failure to fill use "inability to auto-fill". For insufficient filling use "Short Fill". For inconsistent filling - use "Volume Accuracy Problem". | A140302
[6 empty] | Short Fill | Insufficient filling of a device. For complete failure to fill use "inability to auto-fill". For insufficient filling use "short fill". For inconsistent filling - use "Volume Accuracy Problem". | A140303
[6 empty] | Volume Accuracy Problem | Inconsistent filling of a device. This describes a problem which is observed to vary between overfilling and under filling, and may be intermittent. Use "Overfill" or "Short Fill" if problem is consistent. | A140304
|  |  | Filtration Problem | Problem associated with the process of passing a substance through a porous medium, e.g., a blood clot filter for the removal of suspended matter. | A1404 | Inadequate Filtration Process | Problem associated with the filter failing to remove items or substances which should have been removed. | A140401
[6 empty] | Inadequate Ultra Filtration | Problem associated with the transfer of fluid between the blood and dialysate through the dialysis membrane due to a pressure gradient (trans-membrane pressure) existing between the blood and dialysate compartments. | A140402
|  |  | Improper Flow or Infusion | Problem associated with the regulation and delivery of therapeutic agents (e.g. air, gas, drugs or fluids into a device or a patient under positive pressure). | A1405 | Backflow | Continuous flow of fluid (e.g. liquid, gas) against the intended flow direction. | A140501
[6 empty] | Free or Unrestricted Flow | Problem associated with uncontrolled flow of infusion of air, gas or fluids. | A140502
[6 empty] | Gradient Increase | Problem associated with the increased rate of change in temperature, pressure, or other variables as a function of distance, time, etc. | A140503
[6 empty] | Inaccurate Delivery | Delivery at endpoint not as intended; either too low or too high. | A140504
[6 empty] | Inaccurate Flow rate | Problem associated with fluctuations in the flow volume delivered per time, even if end volume is correct, and delivered in the correct total time. | A140505
[6 empty] | Intermittent Infusion | Problem associated with the infusion not being stead- characterised by intermittent stoppages to the flow. | A140506
[6 empty] | Reflux within Device | Problem associated with partial backflow, compromising the device's flow output. | A140507
[6 empty] | Restricted Flow rate | Problem associated with flow rate. Flow volume delivered over time is not reaching intended flow rate. | A140508
[6 empty] | Tidal Volume Fluctuations | Problem associated with the amount of gas that is inspired and expired during one respiratory cycle. | A140509
|  |  | Inflation Problem | Problem associated with the inability of the device to expand or enlarge with the intended inflation agent (e.g. saline or air). | A1406
|  |  | Insufficient Flow or Under Infusion | Problem associated with an insufficient dose of therapeutic agents, e.g., drugs or fluids being delivered into a patient under positive pressure. | A1407
|  |  | No Flow | Problem arising from the device failing to deliver the specified liquid or gas. | A1408 | Failure to Deliver | Failure (=complete nonperformance) with regard to the intended function of delivery. | A140801
[6 empty] | Failure to Infuse | Failure (=complete nonperformance) with regard to the intended function of infusion. | A140802
[6 empty] | Inability to Irrigate | Failure (=complete nonperformance) with regard to the intended function of irrigation. | A140803
|  |  | Obstruction of Flow | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in restriction of flow. | A1409 | Complete Blockage | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in no flow. | A140901
[6 empty] | Partial Blockage | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in a reduction of the flow rate. | A140902
|  |  | Difficult to Flush | The device is difficult to flush, possibly indicating an obstruction within device. | A1410
|  |  | Pressure Problem | Problem associated with the application of a force either internal or external to device that compromises the flow of fluid or gas. | A1411 | Decrease in Pressure | Unintended decrease in pressure, compromising the device's intended function. | A141101
[6 empty] | Increase in Pressure | Unintended increase in pressure, compromising the device's intended function. | A141102
[6 empty] | No Pressure | Unintended complete loss of pressure, compromising the device's intended function. | A141103
|  |  | Pumping Problem | Problem associated with pump performance deviating from specifications in a way to compromise flow or infusion. | A1412 | Decreased Pump Speed | Unintended decrease in pump speed and hence, probably, flow rate, compromising the intended function of the device. | A141201
[6 empty] | Increased Pump Speed | Unintended increase in pump speed and hence, probably, flow rate, compromising the intended function of the device. | A141202
[6 empty] | Failure to Pump | Problem associated with the device which fails to start pumping. | A141203
[6 empty] | Pumping Stopped | Unexpected /unintended cessation of pump. | A141204
|  |  | Suction Problem | Problem associated with suction equipment, which may be a manual, electrical, vacuum or pressure source operated to evacuate and remove undesired substances (air, gas, fluid, or particulates) via tubing and collection bag. | A1413 | Decrease in Suction | Problem associated with the removal of fluid or gas from a body cavity due to decreased suction. | A141301
[6 empty] | Increase in Suction | Problem associated with the removal of excess fluid or gas from a body cavity due to increased suction. | A141302
[6 empty] | Suction Failure | Problem associated with the complete inability to provide suction. | A141303
|  |  | Priming Problem | Problem associated with the preparation of the device to begin pumping. | A1414 | Failure to Prime | Problem associated with the device failing to begin the priming process (i.e. the process of preparation of device for the delivery of fluids). | A141401
[6 empty] | Incomplete or Inadequate Priming | Problem associated with not adequately preparing the device. | A141402
Activation, Positioning or Separation Problem | Problem associated with any deviations from the documented specifications of the device that relate to the sequence of events for activation, positioning or separation of device. NOTE 1 “Deployment” is synonymous with “activation”. | A15 | Activation Problem | Problem associated with the activation of the device. | A1501 | Activation Failure Including Expansion Failures | Problem associated with the device failing to be activated including expansion. | A150101
[6 empty] | Difficult or Delayed Activation | Problem associated with delayed or difficult activation of the device. | A150102
[6 empty] | Premature Activation | Problem associated with early and unexpected activation of the device. | A150103
[6 empty] | Self-Activation or Keying | Problem associated with the unintended activation of the device, or the device having been unexpectedly turned on during use. | A150104
|  |  | Positioning Problem | Problem associated with the movement of the device to an intended location. | A1502 | Positioning Failure | Problem associated with the inability of the device to be positioned in a specified location. | A150201
[6 empty] | Malposition of device | Problem associated with the device being positioned in a location other than intended or specified. | A150202
[6 empty] | Difficult or Delayed Positioning | Problem associated with users experiencing difficulty or delay to position the device to a specified location. | A150203
[6 empty] | Failure to Advance | Problem associated with failure to move the device to an intended location. | A150204
[6 empty] | Difficult to Advance | Problem associated with difficulty moving the device to an intended location (e.g. difficulty in advancing guide wire). | A150205
[6 empty] | Difficult to Insert | Problem associated with problems introducing or inserting the device, even if the user is operating the device in accordance with the instructions for use or labeling. | A150206
[6 empty] | Difficult to Remove | Problem associated with the use of the device in terms of user experiencing difficulty to take out or get rid of the device, even if the user is operating device in accordance with the instructions for use or labeling. | A150207
[6 empty] | Entrapment of Device | Problem associated with the device caught within patient vasculature, tissue, or other device. | A150208
|  |  | Separation Problem | Problem associated with the detachment or separation of the device. | A1503 | Separation Failure | Problem associated with the device or one of its components failing to detach or separate as intended. | A150301
[6 empty] | Difficult or Delayed Separation | Problem associated with users experiencing difficulty or delay with detachment or separation of the device. | A150302
[6 empty] | Premature Separation | Problem associated with an early and unexpected detachment or separation of the device from the system. | A150303
Protective Measures Problem | Problem associated with any deviations from the documented specifications of the device that relate to the implemented and inherited design features specific to devices used for reducing risks to patient or caregiver or maintaining risks within specified levels. | A16 | Device Alarm System | Problem associated with the alarm system of the device. | A1601 | Alarm Not Visible | The device does not display an alarm message when required. | A160101
[6 empty] | No Audible Alarm | The device fails to emit an audible alarm. | A160102
[6 empty] | Low Audible Alarm | The audible device alarm cannot be heard clearly. | A160103
[6 empty] | Delayed Alarm | The device alarm system operates with delay. | A160104
[6 empty] | False Alarm | Problem associated with the device providing incorrect alarm warning or alert to user. | A160105
[6 empty] | Defective Alarm | The device alarm does not operate as expected and/or in agreement with device's specifications. | A160106
|  |  | Fail-Safe Problem | Problem associated with the feature that prevents the unsafe use of the device. | A1602 | Fail-Safe Did Not Operate | Problem associated with the device fail-safe mechanism, which did not function or function in a non effective way, compromising safe use of the device. | A160201
[6 empty] | No Fail-Safe Mechanism | The device does not have a fail-safe mechanism, although such mechanism would be required for its appropriate and/or safe functioning. | A160202
|  |  | Failure of Device to Self-Test | Problem associated with the device failing to perform an internal self-diagnostic process to ensure normal operation during or prior to use. | A1603
|  |  | Failure to Auto Stop | Problem associated with the inability of device to turn itself off when the device is not in an operable condition. | A1604
|  |  | Reset Problem | Problem associated with setting a variable, register, or other storage location back to a prescribed state. | A1605 | Failure to Reset | Problem associated with the device failing to set a variable, register, or other storage location back to a prescribed state. | A160501
[6 empty] | Failure to Zero | Problem associated with the device failing to set a variable, register, or other storage location back to zero. | A160502
[6 empty] | Inappropriate or Unexpected Reset | Problem associated with the device setting a variable, register, or other storage location to an inappropriate or unexpected state. | A160503
|  |  | Premature Indicator Activation | Problems with the activation of a protective measure indicator earlier than expected. | A1606 | Premature Elective Replacement Indicator | Problems with the early or unexpected activation of the elective replacement indicator. | A160601
[6 empty] | Premature End-of-Life Indicator | Problem with the early or unexpected activation of the end-of-life indicator. | A160602
|  |  | Shielding Failure | Problem associated with the device inability to act as a barrier for absorption of radiation energy in X-rays, gamma rays, etc. | A1607
Compatibility Problem | Problem associated with compatibility between device, patients or substances (medication, body fluid, etc.) | A17 | Component or Accessory Incompatibility | Problem associated with the incompatibility of any device while being operated in the same use environment thereby leading to a dysfunction between the devices. | A1701 | Accessory Incompatible | An accessory required for the intended purpose of the device appears incompatible with device, thus compromising the intended function of the device. | A170101
[6 empty] | Component Incompatible | A component required for the proper functioning of the device is not compatible with other components or subassemblies of the device, thus compromising the intended function of the device. | A170102
|  |  | Device-Device Incompatibility | Problem associated with the incompatibility of two or more devices while being operated in the same use environment thereby leading to a dysfunction of more than one device. | A1702
|  |  | Measurement System Incompatibility | Problem associated with the incompatibility of the measurement systems between and/or within device systems that are inherent to the individual device thereby leading to miscalculated or mismatched measurements from those devices, e.g., international metric system versus U.S. measurement system. | A1703
|  |  | Unintended compatibility | Problem associated with the ability of two or more devices which are intended to be incompatible but are able to work or fit together. | A1704
Contamination / decontamination Problem | Problem associated with the presence of any unexpected foreign substance found in the device, on its surface or in the package materials, which may affect performance or intended use of the device, or problem that compromise effective decontamination of the device. | A18 | Contamination During Use | Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device at the user facility. | A1801 | Biofilm coating in Device | Problem associated with the undesired introduction of a biofilm coating into or onto the device. | A180101
[6 empty] | Contamination of Device Ingredient or Reagent | Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device ingredient or reagent. | A180102
[6 empty] | Device Contamination with Body Fluid | Problem associated with the undesired presence of body fluid in/on the device, which are not part of the documented device specifications and requirements. | A180103
[6 empty] | Device Contamination with Chemical or Other Material | Problem associated with contamination of the device with chemical substance or other non biologic material. | A180104
[6 empty] | Microbial Contamination of Device | Problem associated with undesired microbial contamination of the device. | A180105
|  |  | Device Contaminated during manufacture or shipping | Problem associated with the presence of any unexpected foreign substance found on the surface or in the package materials, which may affect optimal performance for its intended use. | A1802
|  |  | Device Reprocessing Problem | Problem associated with a failure during any step of reprocessing process (cleaning, disinfection, packaging, labeling, sterilization) of a used or opened from its original packaging, but unused, device. | A1803 | Failure to Clean Adequately | Problem associated with the failure of the device or operator to remove any visible soil, foreign material or organism deposits on the external surfaces, crevices, and joints of the device. | A180301
[6 empty] | Failure to Disinfect | Failure to properly disinfect the device when reprocessing it. | A180302
[6 empty] | Flushing problem | Flushing process was not executed properly. | A180303
[6 empty] | Problem with Removal of Enzymatic Cleaner | Enzymatic cleaner was not removed properly. | A180304
[6 empty] | Problem with Sterilization | Device was not sterilized properly during reprocessing. | A180305
[6 empty] | Residue After Decontamination | Problem associated with the decontamination process not adequately removing unwanted visible soil, foreign material, or organism deposits. | A180306
Environmental Compatibility Problem | Problem associated with the surrounding conditions in which the device is being used such as temperature, noise, lighting, ventilation, or other external factors such as power supply. | A19 | Ambient Noise Problem | Problem associated with any undesired acoustic energy or vibration that tends to interfere with the operation of the device. | A1901
|  |  | Ambient Temperature Problem | Problem associated with compromised device performance at the ambient temperature or the storage at an inappropriate ambient temperature. | A1902
|  |  | Fumes or Vapors | Problem associated with the visibility, odor, or toxicity of an ambient vapor or gas. | A1903
|  |  | Fungus in Device Environment | Problem associated with the visibility of molds, mildews, yeasts, and/or mushrooms in the immediate environment in which the device is being used. | A1904
|  |  | Moisture or Humidity Problem | Problem associated with an unsatisfactory humidity level in the storage or use environment which affects the device performance. | A1905 | Moisture Damage | Problem associated with damage inflicted upon the device from water vapor or water in the immediate environment in which the device is being used. | A190501
|  |  | Ventilation Problem in Device Environment | Problem associated with the circulation of fresh air in the immediate atmosphere in which the device is being used. | A1906 | Fogging | Problem associated with the visibility of water vapor in the immediate atmosphere in which the device is being used. | A190601
|  |  | Device Unsafe to Use in Environment | Problem associated with environmental condition that results in the unsafe use of the device. (e.g. electromagnetic fields, noise, vibration, microbiological contamination etc.) | A1907
|  |  | Environmental Particulates | Problem associated with fine solids or liquid particles such as dust, smoke, fume, and/or mist suspended in the immediate atmosphere in which the device is being used. | A1908
|  |  | Medical Gas Supply Problem | Problem associated with the facility-supplied medical gases such as medical air, oxygen, nitrous oxide, and nitrogen. | A1909
|  |  | Electrical Power Problem | Problem associated with the quality of the facility-supplied power. | A1910 | Emergency Power Failure | Problem associated with the failure of the facility's emergency power backup system(s) including generators and/or interruptible power systems (UPS). | A191001
[6 empty] | Loss of Power | Problem associated with the failure of primary power supplied by the facility. | A191002
[6 empty] | Power Conditioning Problem | Problem associated with a momentary overpower/over voltage from the utility and electrical systems of user facilities; - Problem associated with inadequate power conditioning such as the presence of fluctuation, surges, spikes, dropouts, noise and other such undesirable transients. | A191003
Installation-Related Problem | Problem associated with unsatisfactory installation, configuration, and/or setup of a specific device. | A20 | Misassembled During Installation | Problem associated with the use of the device characterized by incorrect assembly of device components, parts or constituents. | A2001
Labelling, Instructions for Use or Training Problem | Problem associated with device markings / labelling, instructions for use, training and maintenance documentation or guidelines. | A21 | Device Markings / Labelling Problem | Problem associated with the written, printed or graphic material accompanying or affixed to the device or any of its packaging. This includes verbal instructions relating to identification, technical description, and usage provided by the device manufacturers. Problems can include but are not limited to this material being unclear, missing, worn out, incorrect or inaccurate. | A2101 | Expiration Date Error | Problem associated with errors in identification of expiration date. | A210101
[6 empty] | Illegible Information | Problem associated with information unable to be read or deciphered. | A210102
[6 empty] | Inaccurate Information | Problem associated with imprecise, inexact information. | A210103
[6 empty] | Unclear Information | Problem associated with ambiguous, confused information. | A210104
[6 empty] | Missing Information | Absence of information e.g. labeling, instruction for use. | A210105
|  |  | Lack of Maintenance Documentation or Guidelines | Problem associated with user facility not receiving adequate service documentation, guidelines, or recommendations to perform preventative and corrective maintenance and performance assurance checks. | A2102
|  |  | Inadequate Instructions for Healthcare Professional | Problem associated with inaccuracies in any written, printed, or graphic matter that is affixed to the device or its packaging with any matter that accompanies the device including verbal instructions related to identification, technical description and use of device provided by the device manufacturers that is intended for healthcare professionals. | A2103
|  |  | Inadequate Instructions for Non-Healthcare Professional | Problem associated with users being unclear and not able to follow any written, printed, or graphic matter that is affixed to device or its packaging with any matter that accompanies the device including verbal instructions related to identification, technical description and use of the device provided by the device manufactures that vary from the standard of medical care in a given environment. | A2104
|  |  | Inadequate or Insufficient Training | Problem associated with facility not providing satisfactory initial and/or periodic user training covering operation of the device. | A2105
Human-Device Interface Problem | Problem associated with an act or omission of an act that has a different result than that intended by the manufacturer or expected by the operator. | A22 | Device Difficult to Setup or Prepare | Problem associated with the use of the device in terms of user experiencing difficulty in preparing device for use, even if the operation is being performed according to labeled instructions for use. | A2201
|  |  | Device Difficult to Program or Calibrate | The device is difficult to program, calibrate or set to desired state, even by appropriately trained user/operator. | A2202
|  |  | Device Difficult to Maintain | Problem associated with the user's ability to service the device according to the manufacturer specifications relating to the device routine maintenance, i.e., periodic inspection, failure detection, repair, and care of the device to sustain or restore acceptable operating conditions. | A2203
|  |  | Inadequate User Interface | Problem associated with the means by which the operator and the equipment communicate or interact. | A2204
Use of Device Problem | Problem associated with failure to process, service, or operate the device according to the manufacturer's recommendations or recognized best practices. | A23 | Device Handling Problem | Handling of the device not in accordance with specification, prior to use on the patient. | A2301
|  |  | Use of Incorrect Control Settings | Problem associated with the use of the device in terms of inappropriate and false control setting for the device's specified operation and/or intended use. | A2302
|  |  | Improper or Incorrect Procedure or Method | Problem associated with the use of the device in terms of nonconforming to that device's intended use, specifications, procedure and process or service instructions and information provided by the device manufacturers. | A2303
|  |  | Off-Label Use | Problem associated with the device which has been used for an unapproved indication or for an unapproved intended use. | A2304
|  |  | Misassembled | Problem associated with incorrect assembly of the device or constituents after being put into use. | A2305 | Misassembly by Users | Problem associated with incorrect assembly of the device or constituents by the users. | A230501
[6 empty] | Misassembly during Maintenance / Repair | Problem associated with incorrect assembly of the device or constituents during maintenance or repair. | A230502
Adverse Event Without Identified Device or Use Problem | An adverse event (e.g. patient harm) appears to have occurred, but there does not appear to have been a problem with the device or the way it was used. | A24
No Apparent Adverse Event | A report has been received but the description provided does not appear to relate to an adverse event. This code allows a report to be recorded for administration purposes, even if it doesn't meet the requirements for adverse event reporting. | A25
Insufficient Information | An adverse event appears to have occurred but there is not yet enough information available to classify the device problem. | A26
Appropriate Term/Code Not Available | The device problem is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | A27

## AnnexB Type of Investigation
Annex B: Type of Investigation
Note: Select as many terms as necessary/appropriate to characterise the investigation
Device (bold): For the purpose of this Annex B, a device means a medical device including accessories and components.
Level 1
Term | Definition | Code
Testing of Actual/Suspected Device | The investigation employed relevant empirical testing of the actual device suspected in the reported adverse event in order to establish their functional and other properties and to identify possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B01
Testing of Device from Same Lot/Batch Retained by Manufacturer | The investigation employed relevant empirical testing of the device of the same lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. Testing was performed using the device retained by the manufacturer (i.e. was not shipped). Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B02
Testing of Device from Same Lot/Batch Returned from User | The investigation employed relevant empirical testing of the device of the same lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. The device was returned from the user. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B03
Testing of Device from Other Lot/Batch Retained by Manufacturer | The investigation employed relevant empirical testing of the device of another lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. This includes devices without a lot/batch designation. Testing was performed using the device retained by the manufacturer (i.e. was not shipped). Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B04
Testing of Device from Other Lot/Batch Returned From User | The investigation employed relevant empirical testing of the device of another lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. This includes devices without a lot/batch designation. The device was returned from the user. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the technical file. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B05
Testing of Model Variant | The investigation employed relevant empirical testing of a model variant of the device involved in the reported adverse event in order to support the identification of possible causes for the adverse event through plausibility reasoning. A model variant is not identical to the actual device, but shares relevant characteristics with the device involved. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B06
Testing of Raw/Starting Materials | The investigation employed relevant empirical testing of the materials used in construction of the device involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B07
Testing of Patient Sample or Reference Material Using Manufacturer's Device | The investigation employed relevant empirical testing of a patient sample or reference material using the device (usually an IVD) involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B08
Testing of Patient Sample or Reference Material Using Reference Method | The investigation employed relevant empirical testing of a patient sample or reference material using an appropriate reference method to the device (usually an IVD) involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B09
Testing of Patient Sample or Reference Material Using Competitor's Device | The investigation employed relevant empirical testing of a patient sample or reference material using a competitor's device that is comparable to the device (usually an IVD) involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. | B10
Historical Data Analysis | The investigation involved the analysis of historical adverse events data of the actual device involved in the adverse event and/or of products from the same and/or different batches/lots. | B11
Trend Analysis | The investigation involved trend analysis of adverse event of the actual device involved in the adverse event and/or of products from the same and/or different batches/lots. It should be noted that trend analysis typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. | B12
Communication/Interviews | The investigation involved communication/interviews (either interpersonal or through technical means, e.g. phone, e-mail) with persons close to the adverse event, e.g. healthcare professionals (doctors, nurses etc.), the affected patient(s) or other users including, where appropriate, relatives or others engaged in caring for the affected patient. | B13
Analysis of Production Records | The investigation involved the analysis of relevant production records in view of supporting the identification of possible causes for the adverse event. | B14
Analysis of Data Provided by User/Third Party | The investigation involved the analysis of relevant data provided by the user (e.g. healthcare professional, patient, clinical engineer) or a third party (e.g. testing facility) in view of supporting the identification of possible causes for the adverse event. | B15
Device Not Manufactured by Reporting Manufacturer | Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. | B16
Device Not Returned | The actual device involved in the adverse event was not returned for testing despite requests by manufacturer. | B17
Device Discarded | The actual device involved in the adverse event had been already discarded and thus irretrievably lost for testing. | B18
Incomplete Device Returned | The device was returned incompletely, lacking parts, components or accessories that would be required for appropriate testing and analysis of root causes. | B19
Device Not Accessible for Testing | The actual device involved in the adverse event is not readily accessible for testing (e.g. remains implanted in patient). | B20
Type of Investigation Not Yet Determined | Details to determine the type of investigation are not yet available, but are being sought. Do not use this code if the investigation is complete. | B21
Insufficient Information Available | The information available relating to the reported event is not sufficient to identify either the manufacturer, the device, or other essential information. This term indicates that no further investigation is possible. Do not use this code if further information is being sought, instead use ''Type of investigation not yet determined''. | B22

## AnnexC Investigation Findings
Annex C Investigation Findings ("what were the findings?")
Device (bold): For the purpose of this Annex C, a device means a medical device including accessories and components.
Level 1 |  |  | Level 2 |  |  | Level3
Term | Definition | Code | Term | Definition | Code | Term | Definition | Code
Biological Problem Identified | Problems relating to, caused by or affecting biological processes or living organisms. | C01 | Biocompatibility Problem Identified | The device causes cellular or tissue responses that elicit an undesirable local or systemic effect in the recipient or beneficiary of that therapy. (See ISO 10993) | C0101
|  |  | Biological Contamination | The undesirable presence of living organisms such as bacteria, fungi, or viruses or their products (enzymes or toxins). | C0102 | Endotoxin Contamination | The undesirable presence of toxins associated with certain bacteria (e.g. gram negative bacteria). | C010201
[6 empty] | Microbial Contamination | The undesirable presence of microorganisms or microbes such as bacteria and fungi (yeasts and molds). | C010202
|  |  | Material or Material Leachate Pyrogenic Problem | The undesirable presence of pyrogens or fever-producing organisms caused by materials that permeate through the device. | C0103
|  |  | Cytotoxicity Problem Identified | The device was found to have an undesirable level of toxicity to living cells. | C0104
|  |  | Genotoxicity Problem Identified | The device's ability to cause damage to genetic material (e.g. leading to malignant tumors). (See ISO 10993) | C0105 | Carcinogenic Problem | The device's ability to trigger development of cancer. | C010501
[6 empty] | Mutagenic Problem | The device's ability to change genetic information (usually DNA) of an organism and thus increasing the frequency of mutations. | C010502
|  |  | Hematological Problem Identified | The device affects or impacts the blood or its components. (See ISO 10993 all parts) | C0106 | Agglutination Problem | The device affects the ability of the blood to clot which may be induced by chemical, mechanical, or thermal properties of the device. | C010601
[6 empty] | Complement Activation Problem | The device affects the body's ability to activate the complement system of the immune system, thereby interfering with the ability to clear pathogens. This may be caused by an interaction of the device with chemicals or materials. | C010602
[6 empty] | Platelet Activation Problem | The device affects the body's ability to activate platelet formation. | C010603
[6 empty] | Problem due to Thrombosis Activation | The device causes the formation of blood clots in or along blood vessels resulting in disturbed or disrupted blood flow. | C010604
|  |  | Unintended Presence of Allergens | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labelled, then use "Labelling Problem". | C0107
|  |  | Reproductive Toxicity Problem Identified | The device affects reproductive function, embryo development (teratogenicity), and prenatal and early postnatal development. (ISO 10993 part 3) | C0108
Electrical Problem Identified | Events associated with an electrically powered device where an electrical malfunction results in a device problem (e.g. electrical circuitry, contact or component failed) even if the problem is intermittent. | C02 | Electrical/Electronic Component Problem Identified | The performance of an electrical or electronic component was found to be inadequate. | C0201
|  |  | Hardware Timing Problem Identified | Problems that results from improper sequential activation of components. | C0202
|  |  | Impedance Problem Identified | Problems due to insufficient or excessive resistance to current flow either by the device or circuit. | C0203
|  |  | Insulation Problem Identified | Problems due to inadequate or incorrect electrical insulation material. | C0204
|  |  | Open Circuit | Problem due to an electrical circuit that does not conduct current because a switch is open, a wire is broken, etc. | C0205
|  |  | Current Leakage Problem | Problems related to leakage currents which may cause electric shock. These currents usually flow through the protective ground conductor. In its absence, these currents could flow from the device to the ground via the human body. | C0206
|  |  | Power Source Problem Identified | Problems related to the source that provides electrical power to the device. | C0207 | Energy Storage System Problem | Problems related to the energy storage system (e.g. the rechargeable battery, charging system, or capacitor) and includes problems such as premature power source depletion and battery explosions. | C020701
[6 empty] | Loss of Power | A device that experienced problems due to a loss in the power supply. | C020702
[6 empty] | Power Fluctuation | The device failed due to fluctuations within the power supply (e.g. transient power, power spike, power dip, or power sequencing). | C020703
|  |  | Short Circuit | Problems due to an unintentionally low-resistance connection between two points in an electric circuit, resulting in either excessive current flow that often causes damage or in a new shorter circuit that draws current away from the original pathways and components. | C0208
|  |  | Signal Loss | Problems due to the loss or weakening of an electrical signal or signals. | C0209
Electromagnetic Compatibility Problem Identified | Device-to-device or device-environment problem resulting from electromagnetic disturbances. | C03 | Conducted Interference | Problems related to electromagnetic interference (EMI) by physical contact with conductors (e.g. wires, resistors, terminals) as opposed to radiated EMI which is caused by induction (without physical contact of the conductors). | C0301
|  |  | Electrostatic Discharge | Problems due to sudden and momentary bursts of electrical current flowing between two objects at different electrical potentials. | C0302
|  |  | Inadequate Immunity | Problems related to immunity or capabilities to resist electromagnetic interference (EMI). | C0303
|  |  | Unintended Emission | Problems due to unintended emission of electromagnetic energy by the device. | C0304
|  |  | Radiofrequency Interference (RFI) | Problems due to radiofrequency interference. RFI is a disturbance that affects an electrical circuit due to either electromagnetic conduction or electromagnetic radiation emitted from an external source. | C0305
Interoperability Problem Identified | Problems with the mechanical, electrical, or communication interface between two or more separate devices. | C04 | Communications Problem Identified | Devices that do not send or receive adequate signals (this speaks to the interoperability between devices). | C0401 | Wired Communication Problem | Communications problems between devices within a wired system. | C040101
[6 empty] | Wireless Communication Problem | Communications problems between devices within a wireless system. | C040102
[6 empty] | Network Communication Problem | Communications problems between devices within a network system. | C040103
|  |  | Incompatible Component/ Accessory | A device that malfunctions due to a component(s)/accessory that does not operate correctly and according to the device's specifications. | C0402
|  |  | Device Not Compatible With Another Device | A device that malfunctions due to being used in combination with, or in the presence of, another device. | C0403
|  |  | Unintended Compatibility | The device was confirmed to be compatible with another device with which the device is intended to be incompatible. | C0404
Labeling and Instructions for Use/Maintenance | Insufficient, inadequate, or incorrect information provided on a device's label or documentation regarding e.g. its intended use, directions for use, and characteristics of the device, including its maintenance. | C05 | Inadequate Labelling and/or Instructions for Use | Inadequate information on the labels or in the instructions for use e.g. steps that are difficult to follow or that are missing. | C0501
|  |  | Incorrect Labeling and/or Instructions for Use | Missing, incorrect, or inappropriate information on the labels e.g. mislabeled contents or device labeling characteristics or package contents. | C0502
|  |  | Inadequate or Incorrect Instructions for Maintenance | Inadequate or incorrect information in the instructions for maintenance. | C0503
Material and/or Chemical Problem Identified | Problems with the device materials or how its materials react to other elements either within the device or within the environment. | C06 | Degradation Problem Identified | Problems that occur when the device becomes worn, weakened, corroded, or broken down due to processes such as aging, permeation, and corrosion. | C0601
|  |  | Inappropriate Material | Problems that occur due to the presence of a material that should not be present or part of the device. | C0602 | Improper Composition/ Concentration | Problems associated with the improper combination of materials or elements present in the device (e.g. improper composition of the materials of a capacitor). | C060201
[6 empty] | Improper Physical Structure | Problems related to the incorrect or inadequate arrangement of the parts, components, elements, or materials. | C060202
[6 empty] | Molecular Structure Problem | Problems related to the presence of an inappropriate molecular geometry somewhere in the device (i.e. the spatial arrangement of atoms in a molecule and the chemical bonds that hold the atoms together). | C060203
|  |  | Inadequate Physicochemical Properties | Problems that occur due to the physicochemical properties. | C0603
|  |  | Incompatible Material | Problems that occur due to the incompatibility of materials that co-exist simultaneously as part of the device. | C0604
|  |  | Reactivity Problem Identified | Problems that occur due to the reactivity of materials (e.g. over-react or under-react). | C0605
|  |  | Tolerance Stack-Up | Problems that result from a combination of specification variances of the components. | C0606
Mechanical Problem Identified | Problems that result from internal or external forces including fluids, other objects, or environmental or physiologic influences. | C07 | Device Migration | A device that has moved from its original location due to external forces (e.g. stent or lead movement). | C0701
|  |  | Friction Problem Identified | Problems caused by its surface coming in contact with another surface or fluid. | C0702
|  |  | Leakage/Seal | Problems caused by inadequate/broken seal within the device. | C0703
|  |  | Lubrication Problem Identified | Problems that occurred because of the presence of either too much or too little lubricant where required (e.g. connectors, leading to failure mechanisms such as corrosion). | C0704
|  |  | Stiffness Problem Identified | Problems that occurred when its material is either too flexible/pliable or inflexible/rigid when in contact by an applied force. | C0705
|  |  | Stress Problem Identified | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | C0706 | Deformation Problem | Problems caused by changes in the shape or size of the device due to an applied force. This can be a result of tensile forces, compressive forces, shear, bending, tensile (pulling), or torsion. | C070601
[6 empty] | Fatigue Problem | Problems due to the weakening or breakdown of its material when subjected to stress or a series of repeated stresses. | C070602
[6 empty] | Fracture Problem | Problems caused by the separation of a component, object, or material into two or more pieces including shear. | C070603
[6 empty] | Mechanical Shock Problem | Problems caused by the sudden violent blow or collision to the whole device (e.g. by dropping). | C070604
[6 empty] | Vibration Problem | Problems caused by the constant rhythmic motion of the device, or something in the environment to which the device is exposed. | C070605
[6 empty] | Wear Problem | Problems due to the premature or expected erosion of its material by use, deterioration, or change. | C070606
|  |  | Incorrect Dimension | Problems caused by incorrect physical dimensions of the device or one of its parts | C0707
Optical Problem Identified | Problems related to the optical properties of a device. | C08 | Optical Transmission Problem Identified | Problems with the device's ability to pass light energy. | C0801
|  |  | Light Source Problem Identified | Problems with the optical properties of a device such as diopter, glare, and irradiance or glistening. | C0802
Clinical Imaging Problem Identified | Problems that occur with devices used for radiographic or imaging procedures e.g. CT scanners, magnetic resonance imaging. | C09 | Gradient Induced Field Problem | Problems that result from the gradient-induced fields generated during radiologic procedures e.g. magnetic resonance imaging. | C0901
|  |  | Image Artifact | The unacceptable distortion of an image due to signal loss that may occur during a radiologic procedure such as magnetic resonance imaging. | C0902
|  |  | Magnetically-Induced Movement | Problems due to unintended or excessive movement created by the application of magnetic fields. | C0903
|  |  | Radiofrequency Induced Overheating | Problems due to unintended radiofrequency-induced temperature increase that can occur in the vicinity of the device. | C0904
Software Problem Identified | Problems related to the device software. | C10 | Configuration Issue | Problems due to change control or incorrect version, including regional requirements. | C1001
|  |  | Design Error | The device had faulty (incomplete or incorrect) software design. | C1002 | Data Compression Error | Data was lost or corrupted during the operation of reducing storage space or communication bandwidth. | C100201
[6 empty] | Incorrect Algorithm | The device software was found to implement an incorrect sequence of steps for a specific computation. | C100202
[6 empty] | Incorrect Data Definition | The device software was found to contain errors in specifying or manipulating data items. | C100203
[6 empty] | Interface Design Error | The device software was found to contain errors in the user interface (including usability problems) or the interfaces with other systems. | C100204
[6 empty] | Non-Functional Defect | The device software contained software errors that did not impact its operation. | C100205
[6 empty] | Software Timing Problem | Problems that results from the incorrect sequencing or activation of software modules. | C100206
|  |  | Software Maintenance Problem Identified | The device software was not maintained/updated properly. | C1003
|  |  | Software Installation Problem Identified | The device software was not installed as per the specifications or failed to properly install. | C1004
|  |  | Software Requirement Error | The software requirements for the device are either incomplete, inadequate, or in conflict. | C1005
|  |  | Software Runtime Error | The device software failed during operation as a result of a coding error. | C1006
|  |  | Software Security Vulnerability | The device software failed to provide adequate authorization, access control, protection and accountability features. | C1007
|  |  | Erroneous Data Transfer | The device software fails to transfer the expected data within a system or to another device. | C1008
|  |  | Data Storage or Loss of Data | Storage of data was unsuccessful in total or in part. | C1009
Thermal Problem | Problems related to the temperature of the device. Note: For problems related to environmental temperature use ''Environment Problem Identified''. | C11 | Overheating Problem Identified | The device was found to become hotter than expected during operation. This applies to devices which are not intended to deliver heat. Use "Excessive heating identified" for devices which are intended to deliver heat during operation. Use "Inadequate cooling identified" if the overheating was related to a problem with a cooling system. | C1101
|  |  | Excessive Heating Identified | The device delivered more heat than intended or expected during operation. This applies to devices which are intended to deliver heat. Use "Overheating problem identified" for devices which are not intended to deliver heat during operation. | C1102
|  |  | Excessive Cooling Identified | The device cooled the patient or another device more than intended or expected during operation. | C1103
|  |  | Inadequate Cooling Identified | The device did not sufficiently cool the patient or another device during operation. | C1104
Protective System Problem Identified | Problems related to the system(s) designed to prevent or warn about unsafe operation of the device. | C12 | Fail-safe Problem Identified | A system intended to prevent unsafe operation of the device did not operate correctly. | C1201
|  |  | Alarm System Problem Identified | A system intended to warn of a potentially unsafe condition did not operate correctly. | C1202
|  |  | Problem of Device to Self-Test | Malfunction of the device's self-test system. | C1203
|  |  | Problem to Auto Stop | An auto stop function of a device did not operate correctly. | C1204
|  |  | Premature Indicator Activation | A system intended to indicate the device status was triggered prematurely. | C1205
|  |  | Reset Problem | The device does not reset properly. | C1206
|  |  | Shielding Problem | Inadequate shielding of/by the device. | C1207
|  |  | Missing or Inadequate Safety Measures | Safety measures are inadequately applied or missing. | C1208
Operational Problem Identified | Problems that occur during the performance, use, or functioning of the device. | C13 | Device Incorrectly Reprocessed | Problems associated with the failure to properly and adequately reprocess the device. | C1301 | Device Incorrectly Cleaned During Reprocessing | The cleaning procedure is not followed correctly or used inappropriate cleaning materials. | C130101
[6 empty] | Device Incorrectly Disinfected/Sterilised During Reprocessing | The disinfection/sterilization process was incorrect and/or the wrong products for disinfection/sterilization were used. | C130102
[6 empty] | Device Incorrectly Assembled During Reprocessing | Incorrect assembly of the device following reprocessing. | C130103
|  |  | Failure to Calibrate | A device that cannot calibrate (establish the relationship between a measuring device and the units of measure) to ensure accurate readings. | C1302
|  |  | Device Difficult to Operate | Problems including set-up, operation, and disassembly of equipment. Not including reprocessing. | C1303
|  |  | Incorrect Interpretation of Results/Data | Problems resulting from the incorrect interpretation by the user of the results or data provided by the device. | C1304
Patient Sample Problem | Problems that occurred due to endogenous or exogenous interferent in the sample, or unexpected variation in the target analyte/marker. | C14 | New or Unknown Interferent | New or unknown endogenous or exogenous interferent (sample) identified. | C1401
|  |  | Known Interferent | Known interferent in the sample identified. | C1402
|  |  | Change in Target Marker/Variant/ Mutant | Problem due to change in target marker/variant/mutant which is not covered in the labelling. | C1403
|  |  | Pre-analytical Handling Problem | Incorrect pre-analytical handling of patient's sample by the user. | C1404
Environment Problem Identified | Problems that occurred due to factors within the environment e.g. dust, dirt, humidity, temperature. | C15 | Environmental Temperature Problem Identified | Device performance was affected by the temperature, or changes in temperature, of the environment in which it was used. | C1501
|  |  | Dust or Dirt Problem Identified | A device that experienced problems due to ingress, or coating, of dust or dirt. | C1502
|  |  | Contamination of Environment by Device | Operation of the device results in contamination of the nearby environment e.g. dust, dirt, smoke, heat or biological material. | C1503
|  |  | Environmental Pressure Problem Identified | Device performance was affected by the pressure, or changes in pressure, of the environment in which it was used. | C1504
|  |  | Ambient Light Problem Identified | Device performance was affected by ambient light. This term applies to the direct effects of ambient light on the device, and to the user's ability to operate the device (e.g. to read device output). | C1505
|  |  | Environmental Humidity Problem Identified | Device performance was affected by the humidity, or changes in humidity, of the environment in which it was used. | C1506
Manufacturing Process Problem Identified | Problems with a device that can be traced to a problem in the manufacturing and/or production process. | C16 | Assembly Problem Identified | Problems that occurred because the device was assembled incorrectly. | C1601
|  |  | Sterilization Problem Identified | Problems that occurred during terminal sterilization by the manufacturer. | C1602
|  |  | Installation Problem Identified | A device that malfunctions because it was incorrectly installed, set-up, or configured (e.g. misconfiguration of an "automatic" defibrillator to "semi-automatic", thereby leading to failure). | C1603
|  |  | Maintenance of Manufacturing Machinery | Problems caused by failure to maintain manufacturing equipment used to produce the device. | C1604
|  |  | Packaging Problem Identified | Problems that occurred because of the device packaging. | C1605 | Packaging Compromised | Problems that occurred because of a compromised packaging of the device (e.g. broken or incomplete seal). | C160501
[6 empty] | Packaging Materials Problem | Problems that occurred because of the composition or type of packaging materials was inappropriate for the device. | C160502
[6 empty] | Packaging Contains Unintended Material | Problems that occurred because unintended material was packaged with the device. | C160503
[6 empty] | Packaging Contains Incorrect Device | Problems that occurred because the packaging contained an incorrect device. | C160504
Maintenance Problem Identified | A device malfunction or problem that occurs after production because the device was not properly maintained according to the instructions (e.g. maintenance may be performed by user facility, distributor, or service provider). | C17
Transport/Storage Problem Identified | Problems was caused by transport or storage conditions. | C18 | Transport Problem Identified | Problems traced to how the device was transported e.g. temperature of shipping compartment or method of transportation. | C1801
|  |  | Storage Problem Identified | Problems that result from storing the device in an uncontrolled or improper environment (e.g. moisture sensitive devices stored in a humid environment). | C1802
No Device Problem Found | The device either functioned as intended or a problem was not found. | C19
No Findings Available | Use when no investigation can be performed and therefore no results will be obtained. | C20
Results Pending Completion of Investigation | Investigation is ongoing and results are not yet available. Do not use this code if the investigation is complete. | C21
Appropriate Term/Code Not Available | Problems is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | C22

## AnnexD Investigation Conclusion
Annex D: Investigation Conclusion ("why did the incident/adverse event occur?")
Device (bold): For the purpose of this Annex D, a device means a medical device including accessories and components.
Level 1 |  |  | Level 2
Term | Definition | Code | Term | Definition | Code
Cause Traced to Device Design | Problems traced to the design specifications (e.g. in the requirements, testing processes, hazard analysis, implementation strategy). | D01 | Design Inadequate for Purpose | Problems traced to design/design features of the device that do not support or interfere with the intended purpose of the device. | D0101
|  |  | Human Factors Engineering - Device Difficult to Operate | Problems traced to inappropriate and/or inadequate assessment and engineering design of the device to accommodate how or where the device will be used. | D0102
|  |  | Human Factors Engineering - Device Difficult to Assemble | Problems traced to inadequate design of the component parts and/or assembly steps resulting in the device not being able to be assembled correctly. | D0103
|  |  | Human Factors Engineering - Device Difficult to Reprocess | Problems traced to inadequate design of the reprocessing steps and/or the device resulting in the device remaining unclean. | D0104
|  |  | Missing or Inadequate Safety Measures | Problems traced to inadequate design or complete lack of safety measures leading to device malfunction or unintended properties of the device including possible hazards for persons using the device. | D0105
|  |  | Design Change Validation Inadequate | Problems traced to inadequate or lack of validation of design changes of the device leading to malfunction or unintended properties of the device including possible hazards for persons using the device. | D0106
Cause Traced to Component Failure | Expected or random component failure without any design or manufacturing issue. | D02
Cause Traced to Manufacturing | A defect in the processes or systems used in the manufacture of the device. Examples include problems within the change control, production, or quality control processes. | D03 | Manufacturing Deficiency | Problems traced to manufacturing process. | D0301
|  |  | Quality Control Deficiency | Problems traced to the failure to maintain or establish techniques for controlling and verifying the product specifications (including materials used) identified by the manufacturer himself. | D0302
Cause Traced to Transport/Storage | Problems traced to the inappropriate transport or storage of the device. | D04
Cause Traced to Infrastructure | Problems traced to underlying framework, systems, and processes at the healthcare facility or other point of use (e.g. as building power supply, network, oxygen systems). | D05
Cause Traced to Environment | Problems caused by exposure to environmental conditions outside the expected range. | D06
Cause Traced to Maintenance | Problems traced to improper routine or preventative maintenance. | D07
Cause Traced to Training | Problems caused by inadequate training. | D08
Cause Traced to Labeling | Problems that occur as the result of problems with the labeling (including package inserts, instruction manuals, instructions for use). | D09
Cause cannot be Traced to Device | The adverse event that occurred is not attributable to a device. | D10 | Adverse Event Related to Patient Condition | An existing condition or disease is demonstrably responsible for the adverse event and use of the device has neither caused nor otherwise influenced this condition/disease-related adverse event. | D1001
|  |  | Adverse Event Related to Procedure | The adverse event occurred during the procedure and the device had no influence on event. | D1002
|  |  | Adverse Event Related to Commutability | The adverse event occurred because the material being used to calibrate or assess IVD performance did not have similar properties to those of human samples, leading to inappropriate bias and erroneous results. Examples of materials include: reference materials, calibrators, proficiency testing samples. | D1003
Cause Traced to User | The adverse event caused partially or wholly by the user of the device including sample handling. | D11 | Failure To Follow Instructions | Problems traced to the user not following the manufacturer's instructions. | D1101
|  |  | Unintended Use Error Caused or Contributed to Event | The interaction between the user and device, or sample, caused or contributed to the error. This includes unintended inappropriate use of the device and incorrect sample preparation. | D1102
|  |  | Cause Traced to Intentional Off-Label, Unapproved, or Contraindicated Use | Problems traced to the intentional use of the device in an unapproved procedure, for an unapproved patient, or for which it is contraindicated, or not listed on the label. | D1103
|  |  | Shelf Life/Expiration Date Exceeded | A device that has exceeded the period of time/date recommended by the manufacturer for storing the device without a degradation in quality. | D1104
|  |  | End of Life Problem Identified | Problems traced to the device reaching the end of its useful life. | D1105
|  |  | Reuse of Single Use Device | Problems traced to the use of the device more than once when it is designed for only one use. | D1106
|  |  | Sabotage or Intentional Misuse | Problems occurred because of wilful damage or reckless misuse. This term indicates that there was an apparent intent to cause damage or injury. | D1107
Known Inherent Risk of Device | Reported adverse event known and documented in the labeling (including both short or long term known complications or adverse reactions). | D12
Falsified Device | Devices that deliberately and/or fraudulently misrepresent their identity, composition or source. | D13
No Problem Detected | The device complaint or problem cannot be confirmed. | D14
Cause Not Established | The investigation findings do not lead to a clear conclusion about the cause of the reported adverse event. | D15
Conclusion Not Yet Available | A conclusion has yet to be established as the investigation is incomplete. Do not use this code if the investigation is complete. | D16
Appropriate Term/Code Not Available | The concluded cause is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | D17

## AnnexE Clin signs,sympt,condns
Annex E. Clinical signs, symptoms and conditions
Device (bold): For the purpose of this Annex, a device means a medical device including accessories and components
Wherever appropriate "patient" should be taken to include user, operator or any other person affected by the incident.
LEVEL 1 | LEVEL 2 | [6 empty] | LEVEL 3
Category | Term | Definition | IMDRF Code | MedDRA Code | MedDRA LLT | Primary Category | Secondary Category | Term | Definition | IMDRF Code | MedDRA Code | MedDRA LLT | Primary Category | Secondary Category
Nervous System | Balance Problems | A feeling of falling down which can occur whether the person is standing, sitting or lying down. | E0101 | 10049848 | Balance disorder
| Brain Injury | Damage to the brain. | E0102 | 10060690 | Traumatic brain injury | Nervous System | Injury | Encephalocele | Hernia of brain substance and meninges through a congenital or traumatic opening of the skull. | E010201 | 10014617 | Encephalocele | Nervous System | Injury
| Cerebral Edema | A swelling in the brain caused by the presence of excessive fluid. | E0103 | 10008107 | Cerebral edema | Nervous System | Generalized Disorders
| Cerebral Hyperperfusion Syndrome | Unexpected increase in cerebral blood flow after carotid endarterectomy (CEA) or carotid artery stenting (CAS). | E0104 | 10064730 | Cerebral hyperperfusion syndrome | Nervous System | Vascular System
| Cerebral Ventriculomeglia | Abnormal enlargement of the cerebral ventricles. | E0105 | 10068352 | Cerebral ventriculomegaly
| Cerebrospinal Fluid Leakage | The loss of cerebrospinal fluid into the surrounding tissues. | E0106 | 10008164 | Cerebrospinal fluid leakage
| Cognitive Changes | Changes in perception, thinking, or remembering. | E0107 | 10057668 | Cognitive disorder |  |  | Confusion/ Disorientation | A mental state characterized by a lack of clear and orderly thought and behavior. | E010701 | 10010300 | Confusion
[8 empty] | Dementia | Loss of intellectual abilities interfering with an individual's social and occupational functions. | E010702 | 10012267 | Dementia
| Concussion | Traumatic brain injury as a result of the action of a mechanical force on the head. | E0108 | 10010254 | Concussion | Nervous System | Injury
| Convulsion/Seizure | Sudden, involuntary skeletal muscular contractions of cerebral or brain stem origin. | E0109 | 10010904 | Convulsion |  |  | Convulsion, Clonic | A convulsion marked by alternating contracting and relaxing of the muscles. | E010901 | 10053398 | Clonic convulsion
[8 empty] | Convulsion, Tonic | A convulsion marked by prolonged contraction of the muscles. | E010902 | 10043994 | Tonic convulsion
[8 empty] | Epilepsy | Epilepsy caused or apparently caused by device. Do not use when epilepsy is a pre-existing condition. | E010903 | 10015037 | Epilepsy
[8 empty] | Status Epilepticus | A life-threatening condition caractarized by a single proplonged seizures or a series of seizures without intervening full recovery of conciousness. | E010904 | 10041962 | Status epilepticus
| Decreased Sensitivity | Lower capacity to notice through one or more senses. | E0110 | 10071552 | Hyporesponsive to stimuli
| Increased Sensitivity | Higher capacity to notice through one or more senses. | E0111 | 10082489 | Hyperresponsive to stimuli
| Dizziness | A sensation of lightheadedness, unsteadiness, turning, spinning or rocking. | E0112 | 10013573 | Dizziness
| Dysphasia | Impairment of verbal communication skills, often resulting from brain damage. | E0113 | 10013951 | Dysphasia |  |  | Aphonia | Inability to speak. | E011301 | 10002953 | Aphonia
| Encephalitis | An acute inflammatory process affecting the brain parenchyma. | E0114 | 10014581 | Encephalitis | Nervous System | Infections
| Encephalopathy | A functional and/or structural disorder of the brain that is acquired or congenital and charac terised by mental and neurological symtoms. | E0115 | 10014625 | Encephalopathy
| Headache | Pain in various parts of the head, not confined to the area of distribution of any nerve. | E0116 | 10019211 | Headache | Nervous System | Generalized Disorders
| Hydrocephalus | A disorder characterized by an abnormal increase of cerebrospinal fluid in the ventricles of the brain. | E0117 | 10020508 | Hydrocephalus
| Intracranial Hemorrhage | Bleeding within the cranium. | E0118 | 10022763 | Intracranial hemorrhage
| Loss of consciousness | A level of awareness that can be described as consistently not responsive to stimuli. | E0119 | 10024855 | Loss of consciousness |  |  | Coma | A state of profound unconsciousness associated with markedly depressed cerebral activity. | E011901 | 10010071 | Coma
[8 empty] | Presyncope | Extremely weak; threatened with syncope. | E011902 | 10036653 | Presyncope
[8 empty] | Syncope/Fainting | A spontaneous, transient, loss of consciousness caused by insufficient blood supply to the brain. | E011903 | 10042772 | Syncope
| Memory Loss/Impairment | Systematic and extensive loss or impairment of memory caused by organic or psychological factors. This may be temporary or permanent, and may involve old or recent memories. | E0120 | 10027176 | Memory loss
| Meningitis | A disorder characterized by acute inflammation of the meninges of the brain and/or spinal cord. | E0121 | 10027199 | Meningitis | Nervous System | Infections
| Movement Disorder | Neurological conditions resulting in abnormal voluntary or involuntary movement, which may impact the speed, fluency, quality and ease of movement. | E0122 | 10028035 | Movement disorder |  |  | Dyskinesia | Difficulty moving; distortion or impairment of voluntary movement, as in tic, spasm, or myoclonus. | E012201 | 10013916 | Dyskinesia
[8 empty] | Paralysis | Partial or complete loss of function of one or more muscles. It is usually caused by damage to the nervous system. | E012202 | 10033799 | Paralysis
[8 empty] | Paraplegia | Paralysis of the legs and lower part of the body. | E012203 | 10033892 | Paraplegia
[8 empty] | Paresis | A slight or incomplete paralysis. | E012204 | 10033985 | Paresis
[8 empty] | Quadriplegia | Paralysis of all four limbs. | E012205 | 10037714 | Quadriplegia
[8 empty] | Shaking/Tremors | The shaking movement of the whole body or just a certain part of it, often caused by problems of the neurons responsible for muscle action. | E012206 | 10040528 | Shaking
[8 empty] | Twitching | The occurrence of a single contraction or a series of contractions of a muscle. | E012207 | 10045198 | Twitching
| Nerve Damage | Injury to nervous tissue. | E0123 | 10029177 | Nerve damage | Nervous System | Injury | Optical Nerve Damage | Injury to optical nerve. | E012301 | 10030938 | Optic nerve injury | Nervous System | Injury
| Nervous System Injury | Injury to the central or peripheral nervous system. | E0124 | 10080898 | Nervous system injury | Nervous System | Injury | Spinal Cord Injury | Traumatic damage of the spinal cord. | E012401 | 10041552 | Spinal cord injury | Nervous System | Injury
| Neuralgia | Intense painful sensation along a nerve or group of nerves. | E0125 | 10029223 | Neuralgia
| Neuropathy | A disorder affecting the cranial nerves or the peripheral nervous system. It manifests with pain, tingling, numbness, and muscle weakness. | E0126 | 10029328 | Neuropathy
| Numbness | The loss of the sensation of feeling in an area of the body. | E0127 | 10029829 | Numbness
| Peripheral Nervous Injury | Injury to the peripheral nervous system. | E0128 | 10034586 | Peripheral nerve injury | Nervous System | Injury
| Sensitivity of Teeth | Pain or discomfort caused by exposure of the dentin layer of tooth to thermal, tactile, or other stimuli. | E0129 | 10040012 | Sensitivity of teeth
| Sleep Dysfunction | A change from the patient's baseline sleeping pattern, in the hours slept and/or an alteration/dysfunction in the stages of sleep. | E0130 | 10040984 | Sleep disorder
| Speech Disorder | A term referring to disorders characterized by the disruption of normal speech. It includes stuttering, lisps, dysarthria and voice disorders. | E0131 | 10041466 | Speech disorder
| Spinal Arachnoiditis | A chronic adhesive arachnoiditis in the spinal arachnoid, with root and spinal cord symptoms similar to those caused by pressure from a tumor. | E0132 | 10068160 | Spinal arachnoiditis
| Stroke/CVA | A sudden loss of neurological function secondary to hemorrhage or ischemia in the brain parenchyma due to a vascular event. | E0133 | 10042244 | Stroke |  |  | Hemorrhagic Stroke | An acute neurological deficit caused by bleeding within the brain or between the brain and the skull. | E013301 | 10048863 | Hemorrhagic stroke
[8 empty] | Ischemia Stroke | An ischemic condition of the brain, producing a persistent focal neurological deficit in the area of distribution of the cerebral arteries. | E013302 | 10055221 | Ischemic stroke
| Tactile Disorders | conditions characterized by an alteration in the ability to or perception of touch. | E0134 | 10062872 | Dysesthesia |  |  | Hyperesthesia | Increased sensitivity to tactile stimulation. | E013401 | 10020622 | Hyperesthesia
[8 empty] | Hypoesthesia | Impairment of tactile sensitivity manifesting as partial loss of sensitivity to sensory stimuli. | E013402 | 10020976 | Hypoesthesia
[8 empty] | Paresthesia | Abnormal touch sensations, such as burning or prickling, that occur without an outside stimulus. | E013403 | 10033987 | Paresthesia
| Taste Disorder | Definition:conditions characterized by an alteration in taste function or perception | E0135 | 10082490 | Taste disorder |  |  | Ageusia | Inability to taste. | E013501 | 10001480 | Ageusia
[8 empty] | Dysgeusia | Distortion of taste. | E013502 | 10013911 | Dysgeusia
| Tics/Tremor | An involuntary, compulsive, repetitive stereotyped movement. | E0136 | 10066770 | Motor tic
| Transient Ischemic Attack | A brief attack (from a few minutes to an hour) of cerebral dysfunction of vascular origin, with no persistent neurological deficit. | E0137 | 10072760 | Transient ischemic attack
| Undesired Nerve Stimulation | Unintended reaction in nerve system caused by electrical stimulation. | E0138 | 10068145 | Undesired device stimulation of nerve
| Acoustic Shock | Neurophysiological and psychological symptoms caused by sudden and unexpected loud noise. | E0901 | 10068150 | Acoustic shock | Ear and Labyrinth | Nervous System
| Muscle Weakness | A reduction in the strength of one or more muscles. | E1621 | 10028350 | Muscle weakness | Musculoskeletal System | Nervous System
| Perinatal Brain Injury | Injury to the brain of a newborn infant occurring around the time of birth. | E1513 | 10056392 | Perinatal brain damage | Pregnancy, Childbirth and the Puerperium | Nervous System
| Tinnitus | A noise in the ears, such as ringing, buzzing, roaring or clicking in the absence of auditory stimulation. | E0904 | 10043882 | Tinnitus | Ear and labyrinth | Nervous System
Mental, Emotional and Behavioural Disorders | Bruxism | Excessive clenching of the jaw and grinding of the teeth. | E0201 | 10006514 | Bruxism
| Emotional Changes | Change from the affected person's usual emotional state. | E0202 | 10014551 | Emotional disorder |  |  | Anxiety | Apprehension or uneasiness which may be accompanied by restlessness, tension, tachycardia, and dyspnea. | E020201 | 10002855 | Anxiety
[8 empty] | Depression | A condition marked by ongoing feelings of sadness, despair, loss of energy, diminished interest/pleasure and difficulty dealing with normal daily life. | E020202 | 10012378 | Depression
[8 empty] | Distress | A state of physiological or psychological stress that cannot be compensated for by normal adaptive measures. | E020203 | 10049119 | Emotional distress
[8 empty] | Malaise | A feeling of general discomfort or uneasiness, an out-of-sorts feeling. | E020204 | 10025482 | Malaise
[8 empty] | Irritability | A feeling of agitation. Tending to become frustrated or upset easily. | E020205 | 10022998 | Irritability
| Hallucination | A false sensory perception in the absence of an external stimulus. | E0203 | 10019063 | Hallucination
| Lethargy | Characterized by a lack of vitality or energy. | E0204 | 10024264 | Lethargy
| Suicidal Ideation | Thoughts of taking one's own life. | E0205 | 10042458 | Suicidal ideation
Blood and Lymphatic System | Anemia | A reduction in the number of red blood cells, the amount of hemoglobin, and/or the volume of packed red blood cells. | E0301 | 10002272 | Anemia |  |  | Hemolytic Anemia | Anemia resulting from the premature destruction of the peripheral blood red cells. | E030101 | 10019493 | Hemolytic anemia
| Coagulation Disorder | A condition in which there is a deviation from or interruption of the normal coagulation properties of the blood. | E0302 | 10009731 | Coagulation disorder |  |  | Disseminated Intravascular Coagulation (DIC) | A pathological process where the blood starts to coagulate throughout the whole body. This depletes the body of its platelets and coagulation factors, and there is an increased risk of hemorrhage. | E030201 | 10013442 | Disseminated intravascular coagulation
[8 empty] | Thrombocytopenia | Low blood platelet count. | E030202 | 10043554 | Thrombocytopenia
| Hemolysis | Disruption of the integrity of the erythrocyte membrane causing release of hemoglobin. | E0303 | 10019491 | Hemolysis
| Hypervolemia | Too much fluid in the blood. | E0304 | 10020920 | Hypervolemia
| Hypovolemia | An abnormally low circulating blood volume. | E0305 | 10021139 | Hypovolemia
| Sepsis | The presence of pathogenic microorganisms in the blood stream causing a rapidly progressing systemic reaction that may lead to shock. Symptoms include fever, chills, tachycardia, and increased respiratory rate. It is a medical emergency that requires urgent medical attention. | E0306 | 10040047 | Sepsis | Blood and Lymphatic System | Infections
| Seroma | A tumor-like collection of serum in the tissues. | E0307 | 10040102 | Seroma
| Swollen Lymph Nodes/Glands | The presence of an enlarged lymph node. | E0308 | 10067201 | Swollen lymph nodes
| Electrolyte Imbalance | Higher or lower than normal values for the serum electrolytes; usually affecting Na, K, Cl, Ca, CO2, BUN. | E1202 | 10014418 | Electrolyte imbalance | Metabolism, and Nutrition | Blood and Lymphatic System | Hypernatremia | Higher than normal levels of sodium in the circulating blood. | E120201 | 10020680 | Hypernatremia | Metabolism, and Nutrition | Blood and Lymphatic System
[8 empty] | Hyponatremia | Lower than normal levels of sodium in the circulating blood. | E120202 | 10021038 | Hyponatremia | Metabolism, and Nutrition | Blood and Lymphatic System
Immune System | Autoimmune Disorder | A disorder resulting from loss of function or tissue destruction of an organ or multiple organs, arising from humoral or cellular immune responses of the individual to his own tissue constituents. | E0401 | 10061664 | Autoimmune disorder
| Hypersensitivity/Allergic reaction | A local or general reaction of an organism following contact with a specific allergen to which it has been previously exposed and to which it has become sensitized. | E0402 | 10020751 | Hypersensitivity |  |  | Anaphylactic Shock | Progression of anaphylaxis into systemic shock, which is characterized by significant reduction in tissue perfusion which leads to hypoxia and hypotension, and can lead to death if untreated. | E040201 | 10002199 | Anaphylactic shock | Immune System | Generalized Disorders
[8 empty] | Anaphylactoid | Resembling anaphylaxis | E040202 | 10002216 | Anaphylactoid reaction | Immune System | Generalized Disorders
[8 empty] | Local Reaction | A regionally-limited response to an antigen, which may include inflammation, induration, erythema, pruritus or pain. | E040203 | 10024769 | Local reaction
| Immunodeficiency | A deficiency of immune response or a disorder characterized by deficient immune response. | E0403 | 10061598 | Immunodeficiency
| Positive antinuclear antibodies (ANA) | Indicates that antibodies directed against nuclear antigens have been detected. Almost invariably found in systemic lupus erythematosus and are frequently found in rheumatoid arthritis, scleroderma, Sjogren's Syndrome and mixed connective tissue disease. | E0404 | 10060055 | Antinuclear antibody positive
Vascular System | Aneurysm | Bulging or ballooning in an area of an artery secondary to arterial wall weakening. | E0501 | 10002329 | Aneurysm |  |  | Ruptured Aneurysm | An aneurysm that has burst, resulting in hemorrhage into adjacent tissues. | E050101 | 10048380 | Aneurysm ruptured
| Arteriosclerosis/ Atherosclerosis | A vascular disorder characterized by thickening and hardening of the walls of the arteries. Build-up of fatty material and calcium deposition in the arterial wall resulting in partial or complete occlusion of the arterial lumen. | E0502 | 10003210 | Arteriosclerosis
| Embolism/Embolus | The blockage of a blood vessel lumen by air or solid material such as blood clot or other tissues (e.g., adipose tissue, cancer cells) that have migrated from another anatomic site. | E0503 | 10061169 | Embolism |  |  | Air Embolism | The presence of bubbles of air in the vascular system; occurrence is related to the entry of air into the venous circulation following trauma or surgery. | E050301 | 10001526 | Air embolism
[8 empty] | Foreign Body Embolism | Type of embolism when lumen of vessel or ductus is occluded by part of device. | E050302 | 10075247 | Foreign body embolism
[8 empty] | Pulmonary Embolism | The obstruction of the pulmonary artery or one of its branches by an embolus, sometimes associated with infarction of the lung. | E050303 | 10037377 | Pulmonary embolism | Vascular System | Respiratory System
[8 empty] | Thromboembolism | Occlusion of the lumen of a vessel by a thrombus that has migrated from a distal site via the blood stream. | E050304 | 10043566 | Thromboembolism
| Extravasation | The leakage of body fluid or cells from the vascular system to the surrounding tissues. This process occurs either by force or as a result of a pathologic reaction (e.g., inflammation). | E0504 | 10015866 | Extravasation
| Hematoma | A localized collection of blood, usually clotted, in an organ, space, or tissue, due to a break in the wall of a blood vessel. | E0505 | 10019428 | Hematoma
| Hemorrhage/Bleeding | The loss of blood from a ruptured blood vessel. | E0506 | 10019524 | Hemorrhage |  |  | Ecchymosis | A flat discoloration of the skin or a mucous membrane, larger in area than a petechia, caused by the extravasation of blood. Discolouration under the skin caused by a bleed underneath. Typically a bruise. | E050601 | 10014080 | Ecchymosis
[8 empty] | Exsanguination | Extensive loss of blood due to internal or external hemorrhage. | E050602 | 10015719 | Exsanguination
| High Pulmonary Arterial Wedge Pressure | The measurement of the mean left atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is high. | E0507 | 10037328 | Pulmonary arterial wedge pressure high | Vascular System | Investigations and Diagnostic Tests
| Low Pulmonary Arterial Wedge Pressure | The measurement of the mean left pulmonary atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low. | E0508 | 10037330 | Pulmonary arterial wedge pressure low | Vascular System | Investigations and Diagnostic Tests
| Ischemia | Lack of oxygen in tissues, commonly due to impaired blood supply to an area of the body, resulting in impairment of tissue oxygenation. | E0509 | 10054438 | Ischemia
| Neovascularization | The formation of new blood vessels. | E0510 | 10029114 | Neovascularization
| Perforation of Vessels | Damage to a blood vessel resulting from a full thickness disruption in the integrity of the vessel wall. | E0511 | 10068149 | Vessel perforation | Vascular System | Procedural Complications | Great Vessel Perforation | Damage to the venae cavae, pulmonary artery, pulmonary veins or aorta resulting from a full thickness disruption in the integrity of the vessel wall. | E051101 | 10080896 | Perforation of great vessels | Vascular System | Procedural Complications
| Phlebitis | Inflammation of a vein. | E0512 | 10034879 | Phlebitis
| Pseudoaneurysm | The creation of hematoma outside the arterial wall following bleeding due to puncture of the arterial wall. Pseudoaneurysms can also occur in the heart chambers following myocardial infarction with rupture of the wall of the ventricle. | E0513 | 10048977 | Pseudoaneurysm
| Thrombosis/Thrombus | The formation of a blood clot in the lumen of a vessel or heart chamber. | E0514 | 10043607 | Thrombosis
| Vascular Dissection | Spontaneous or induced tear within the wall of a blood vessel. | E0515 | 10070693 | Vascular dissection
| Vasoconstriction | Vasoconstriction involves narrowing of a blood vessel lumen due to smooth muscle contraction in the vessel wall and leads to decreased blood flow. | E0516 | 10047139 | Vasoconstriction
| Vasodilatation | Widening of a blood vessel lumen due to smooth muscle relaxation in the vessel wall and leads to increased blood flow. | E0517 | 10047141 | Vasodilatation
| Bronchial Hemorrhage | Bleeding from the bronchial wall and/or lung parenchyma. | E0707 | 10065757 | Bronchial hemorrhage | Respiratory System | Vascular System
| Cardiovascular Insufficiency | Circulatory insufficiency due to a decrease in the force of the cardiac contractions and / or the tone of the vascular walls | E0607 | 10065929 | Cardiovascular insufficiency | Heart | Vascular System
| Cerebral Hyperperfusion Syndrome | Unexpected increase in cerebral blood flow after carotid endarterectomy (CEA) or carotid artery stenting (CAS). | E0104 | 10064730 | Cerebral hyperperfusion syndrome | Nervous System | Vascular System
Heart | Arrhythmia | Any variation from the normal rate or rhythm (which may include the origin of the impulse and/or its subsequent propagation) in the heart. | E0601 | 10003119 | Arrhythmia |  |  | Asystole | An electrocardiographic finding showing a state of cardiac standstill with no cardiac output and no ventricular depolarization. | E060101 | 10003586 | Asystole
[8 empty] | Atrial Fibrillation | Irregular heart rhythm, usually associated with a rapid rate, that is caused by abnormal electrical activity within the atrial. | E060102 | 10003658 | Atrial fibrillation
[8 empty] | Atrial Flutter | A rapid heart rate, associated with a regular rhythm, that is caused by abnormal electrical activity within the atria. | E060103 | 10003662 | Atrial flutter
[8 empty] | Bradycardia | An abnormally slow heart rate. Thresholds for different age, gender, and patient populations exist. | E060104 | 10006093 | Bradycardia
[8 empty] | Ectopic Heartbeat | A cardiac beat originating from an abnormal site. These changes lead to extra or dropped heartbeats. | E060105 | 10014159 | Ectopic heartbeats
[8 empty] | Heart Block | An interference with the normal conduction of electrical impulses that control activity of the heart muscle. | E060106 | 10019252 | Heart block
[8 empty] | Idioventricular Rhythm | An electrocardiographic finding of three or more consecutive complexes of ventricular origin. The QRS complexes are wide and have an abnormal morphology. (CDISC) | E060107 | 10039111 | Rhythm idioventricular
[8 empty] | Irregular Pulse | A pulse that beats at irregular intervals. | E060108 | 10022994 | Irregular pulse
[8 empty] | Tachycardia | Abnormally high heart rate. Thresholds for different age, gender, and patient populations exist. | E060109 | 10043071 | Tachycardia
[8 empty] | Ventricular Fibrillation | Arrhythmia characterized by chaotic contractions of the ventricular muscle. | E060110 | 10047290 | Ventricular fibrillation
| Cardiac Arrest | The sudden cessation of productive cardiac activity in an individual who becomes unresponsive, without normal breathing and no signs of circulation. | E0602 | 10007515 | Cardiac arrest
| Cardiac Enzyme Elevation | Increase from the normal levels of cardiac enzymes. | E0603 | 10007548 | Cardiac enzymes increased | Heart | Investigations and Diagnostic Tests
| Cardiac Perforation | The presence of an acquired hole in the heart. | E0604 | 10058039 | Cardiac perforation | Heart | Procedural Complications
| Cardiac Tamponade | Acute compression of the heart caused by increased intrapericardial pressure due to the accumulation of blood or fluid in the pericardium from rupture of the heart or great vessels, penetrating trauma, or progressive effusion. | E0605 | 10007610 | Cardiac tamponade
| Cardiomyopathy | A disease of the heart muscle. | E0606 | 10007636 | Cardiomyopathy
| Cardiovascular Insufficiency | circulatory insufficiency due to a decrease in the force of the cardiac contractions and / or the tone of the vascular walls | E0607 | 10065929 | Cardiovascular insufficiency | Heart | Vascular System
| Cusp Tear | A tear located on a leaflet of a heart valve that prevents its complete closing, which can result in valvular dysfunction. | E0608 | 10068167 | Cusp tear
| Diminished Pulse Pressure | Weak or absent palpable pulse due to decreased arterial pulse pressure. | E0609 | 10037480 | Pulse pressure decreased
| Endocarditis | Inflammation of the endocardium. | E0610 | 10014665 | Endocarditis | Heart | Infections
| Heart Failure/Congestive Heart Failure | Inability of the heart to pump blood adequately to fill tissue metabolic requirements or the ability to do so only at an elevated filling pressure. | E0611 | 10019279 | Heart failure
| Ischemic Heart Disease | A disorder of cardiac function caused by insufficient blood flow to the muscle tissue of the heart. The decreased blood flow may be due to narrowing of the coronary arteries, to obstruction by a thrombus, or less commonly, to diffuse narrowing of arterioles and other small vessels within the heart. Severe interruption of the blood supply to the myocardial tissue may result in necrosis of cardiac muscle (myocardial infarction). | E0612 | 10055218 | Ischemic heart disease |  |  | Angina | Chest pain resulting from inadequate oxygen delivery for the needs of the myocardium. | E061201 | 10002383 | Angina pectoris
[8 empty] | Myocardial Infarction | A condition which occurs when blood flow decreases or stops to a part of the heart, causing damage to the heart muscle. | E061202 | 10028596 | Myocardial infarction
| Low Cardiac Output | A below normal volume of blood pumped from the right or left ventricle of the heart per unit time. | E0613 | 10007595 | Cardiac output decreased
| Mitral Perforation | A rupture in the mitral valve tissue due to traumatic or pathologic processes. | E0614 | 10068138 | Mitral perforation | Heart | Procedural Complications
| Myocardial Contusion | A bruise to the heart. | E0615 | 10068147 | Myocardial contusion | Heart | Procedural Complications
| Myocardial Hypertrophy | Thickening of the myocardium often due to chronic pressure overload. | E0616 | 10072733 | Myocardial hypertrophy
| Myocarditis | Inflammation of the muscle tissue of the heart. | E0617 | 10028606 | Myocarditis
| Non specific EKG/ECG Changes | An electrocardiographic finding of a change in cardiac electrical activity. | E0618 | 10061116 | Electrocardiogram change
| Pericardial Effusion | Fluid collection within the pericardial sac, usually due to inflammation. | E0619 | 10061116 | Electrocardiogram change
| Pericarditis | Inflammation of the pericardium. | E0620 | 10034484 | Pericarditis
| Valvular Insufficiency/ Regurgitation | Dysfunction of one of the cardiac valves, with incomplete valve closure or damage to a valve leaflet resulting in valvular backflow. | E0621 | 10058086 | Heart valve insufficiency |  |  | Aortic Valve Insufficiency/ Regurgitation | Incomplete closure or damage to a valve leaflet of the aortic valve, resulting in aortic valve backflow. | E062101 | 10052839 | Aortic valve insufficiency
[8 empty] | Mitral Valve Insufficiency/ Regurgitation | Incomplete closure or damage to a valve leaflet of the mitral valve, resulting in mitral valve backflow. | E062102 | 10074862 | Mitral valve insufficiency
[8 empty] | Pulmonary Valve Insufficiency/ Regurgitation | Incomplete closure or damage to a valve leaflet of the pulmonary valve, resulting in pulmonary valve backflow. | E062103 | 10052842 | Pulmonary valve insufficiency
[8 empty] | Tricuspid Valve Insufficiency/ Regurgitation | Incomplete closure or damage to a valve leaflet of the tricuspid valve, resulting in tricuspid valve backflow. | E062104 | 10044635 | Tricuspid insufficiency
| Valvular Stenosis | Obstruction of flow across a heart valve due to immobility of the leaflets and/or a restricted orifice. | E0622 | 10061996 | Heart valve stenosis |  |  | Aortic Valve Stenosis | Obstruction of flow across a aortic heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | E062201 | 10002918 | Aortic valve stenosis
[8 empty] | Mitral Valve Stenosis | Obstruction of flow across a mitral heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | E062202 | 10027733 | Mitral valve stenosis
[8 empty] | Pulmonary Valve Stenosis | Obstruction of flow across a pulmonary heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | E062203 | 10037450 | Pulmonary valve stenosis
[8 empty] | Tricuspid Valve Stenosis | Obstruction of flow across a tricuspid heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | E062204 | 10044642 | Tricuspid valve stenosis
| Fetal Distress | signs or symptoms which may be indicative of distress to a fetus, which may include repetitive variable decelerations, fetal tachycardia or bradycardia, late decelerations, or low biophysical profile. (new definition) | E1507 | 10016486 | Fetal distress | Pregnancy, Childbirth and the Puerperium | Heart
Respiratory System | Adult Respiratory Distress Syndrome | Progressive and life-threatening pulmonary distress in the absence of an underlying pulmonary condition, usually following major trauma or surgery. Cases of neonatal respiratory distress syndrome are not included in this definition. | E0701 | 10001409 | Adult respiratory distress syndrome
| Airway Obstruction | Blockage of air flow in the respiratory tract. | E0702 | 10001539 | Airways obstruction
| Apnea | Transient cessation of respiration. | E0703 | 10002972 | Apnea
| Aspiration/Inhalation | The entry of foreign material past the trachea, and into the lower respiratory tract. | E0704 | 10003504 | Aspiration
| Aspiration Pneumonitis | Inflammation of the lungs due to the inhalation of solid or liquid material. | E0705 | 10003526 | Aspiration pneumonitis
| Asthma | A chronic respiratory disease manifested as difficulty breathing due to the narrowing of bronchial passageways. | E0706 | 10003553 | Asthma
| Bronchial Hemorrhage | Bleeding from the bronchial wall and/or lung parenchyma. | E0707 | 10065757 | Bronchial hemorrhage | Respiratory System | Vascular System
| Bronchitis | An acute or chronic inflammatory process affecting the bronchi. | E0708 | 10006451 | Bronchitis
| Bronchospasm | Sudden contraction of the smooth muscles of the bronchial wall. | E0709 | 10006482 | Bronchospasm
| Choking | Interference with respritation by ostruction of the larynx or trachea. | E0710 | 10008589 | Choking | Respiratory System | Injury
| Chronic Obstructive Pulmonary Disease (COPD) | A chronic and progressive lung disorder characterized by the loss of elasticity of the bronchial tree and the air sacs, destruction of the air sacs wall, thickening of the bronchial wall, and mucous accumulation in the bronchial tree. The pathologic changes result in the disruption of the air flow in the bronchial airways. Signs and symptoms include shortness of breath, wheezing, productive cough, and chest tightness. The two main types of chronic obstructive pulmonary disease are chronic obstructive bronchitis and emphysema. | E0711 | 10009033 | Chronic obstructive pulmonary disease |  |  | Pulmonary Emphysema | A condition characterized by increase beyond normal in the size of air spaces distal to the terminal bronchioles, either from dilatation of the alveoli or from destruction of their walls. | E071101 | 10014563 | Emphysema pulmonary
| Cough | A sudden, often repetitive, spasmodic contraction of the thoracic cavity, resulting in violent release of air from the lungs, and usually accompanied by a distinctive sound. | E0712 | 10011224 | Cough
| Decreased Peak Expiratory Flow rate | Decrease in the maximum flow rate generated during a forceful exhalation, starting from full lung inflation. | E0713 | 10034195 | Peak expiratory flow rate decreased | Respiratory System | Investigations and Diagnostic Tests
| Increased Peak Expiratory Flow rate | An increase of the greatest rate of flow that can be achieved during forced expiration beginning with the lungs fully inflated. | E0714 | 10034196 | Peak expiratory flow rate increased | Respiratory System | Investigations and Diagnostic Tests
| Decreased Respiratory Rate | Abnormal decrease of rate of breathing. | E0715 | 10038710 | Respiratory rate decreased | Respiratory System | Investigations and Diagnostic Tests
| Increased Respiratory Rate | Abnormal increase of rate of breathing. | E0716 | 10038712 | Respiratory rate increased | Respiratory System | Investigations and Diagnostic Tests
| Dyspnea | An uncomfortable sensation of difficulty breathing. It may present as an acute or chronic sign of an underlying respiratory or heart disorder. | E0717 | 10013963 | Dyspnea
| Epistaxis | Nose bleed. | E0718 | 10015090 | Epistaxis
| Forced Expiratory Volume Decreased | A decrease in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | E0719 | 10016987 | Forced expiratory volume decreased | Respiratory System | Investigations and Diagnostic Tests
| Forced Expiratory Volume Increased | An increase in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | E0720 | 10016989 | Forced expiratory volume increased | Respiratory System | Investigations and Diagnostic Tests
| Hemoptysis | Coughing up blood from the respiratory tract. | E0721 | 10019523 | Hemoptysis
| Hemothorax | A collection of blood in the pleural cavity. | E0722 | 10019614 | Hemothorax
| Hiccups | An involuntary contraction of the diaphragm against closed vocal cords, producing the "hic" sound. | E0723 | 10020039 | Hiccups
| Hyperventilation | Abnormally prolonged, rapid, and deep breathing. | E0724 | 10020910 | Hyperventilation
| Hypoventilation | A state in which there is a reduced amount of air entering the pulmonary alveoli. | E0725 | 10021133 | Hypoventilation
| Hypoxia | A deficiency in the amount of oxygen reaching the tissues. | E0726 | 10021143 | Hypoxia
| Mediastinal Shift | Displacement of structures in the thoracic cavity in one direction due to increased pleural pressure on the contralateral side. | E0727 | 10066239 | Mediastinal shift
| Nasal Obstruction | Blockage of the nasal passages. | E0728 | 10028748 | Nasal obstruction
| Overinflation of Lung | Abnormal permanent enlargement of the lung air spaces distal to terminal bronchiole not resulted from wall destruction, e.g. due to loss of opposite lung. | E0729 | 10059487 | Lung hyperinflation
| Pharyngitis | Inflammation of the throat most often caused by viral and bacterial infections. | E0730 | 10034835 | Pharyngitis
| Pleural Effusion | Increased amounts of fluid within the pleural cavity. Symptoms include shortness of breath, cough, and chest pain. | E0731 | 10035598 | Pleural effusion
| Pleural Empyema | An accumulation of pus in the pleural space | E0732 | 10037381 | Pulmonary empyema
| Pneumonia | An acute and/or chronic inflammation focally or diffusely affecting the lung parenchyma. | E0733 | 10035664 | Pneumonia | Respiratory System | Infections | Bronchopneumonia | Acute inflammation of the walls of the terminal bronchioles. | E073301 | 10006469 | Bronchopneumonia | Respiratory System | Infections
| Pneumothorax | Abnormal presence of air in the pleural cavity. | E0734 | 10035759 | Pneumothorax
| Pulmonary Dysfunction | Impaired or abnormal functioning of the pulmonary system causing multiple pulmonary issues. | E0735 | 10037386 | Pulmonary function impairment
| Pulmonary Edema | Accumulation of fluid in the lung tissues causing disturbance of the gas exchange that may lead to respiratory failure. | E0736 | 10037375 | Pulmonary edema | Respiratory System | Generalized Disorders
| Pulmonary Hypertension | Increased pressure within the pulmonary circulation due to lung or heart disorder. | E0737 | 10037400 | Pulmonary hypertension
| Pulmonary Infarction | Localized necrosis of lung tissue caused by obstruction of the arterial blood supply, most often due to pulmonary embolism. | E0738 | 10037410 | Pulmonary infarction
| Respiratory Acidosis | A condition in which the blood pH is less than normal, secondary to impaired gas exchange. | E0739 | 10038661 | Respiratory acidosis | Respiratory System | Metabolism, and Nutrition
| Respiratory Alkalosis | A condition in which the blood pH is greater than normal, secondary to impaired gas exchange. | E0740 | 10038664 | Respiratory alkalosis | Respiratory System | Metabolism, and Nutrition
| Respiratory Arrest | Cessation of breathing function. | E0741 | 10038669 | Respiratory arrest
| Respiratory Failure | The significant impairment of gas exchange within the lungs resulting in hypoxia, hypercarbia, or both, to the extent that organ tissue perfusion is severely compromised. | E0742 | 10038695 | Respiratory failure
| Respiratory Insufficiency | Impairment of gas exchange within the lungs secondary to a disease process, neoplasm, or trauma, possibly resulting in hypoxia, hypercarbia, or both. | E0743 | 10038701 | Respiratory insufficiency
| Respiratory Tract Infection | An infectious disease affecting the respiratory tract. | E0744 | 10062352 | Respiratory tract infection | Respiratory System | Infections
| Sinus Perforation | A rupture in the sinus tissue due to traumatic or pathologic processes. | E0745 | 10040748 | Sinus perforation | Respiratory System | Procedural Complications
| Sneezing | A symptom consisting of the involuntary expulsion of air from the nose. | E0746 | 10041232 | Sneezing
| Sore Throat | Any kind of inflammatory process of the tonsils, pharynx, or/and larynx characterized by pain in swallowing. | E0747 | 10041367 | Sore throat
| Stacking Breaths | Condition where there are a greater number of inhalations relative to the number of exhalations. | E0748 | 10068523 | Breath stacking
| Strangulation | Interference with respiration by compression of the larynx, laryngopharynx or trachea,as well as of the carotid arteria or jugular veins. | E0749 | 10042169 | Strangulation | Respiratory System | Injury
| Ventilator Dependent | A condition that requires mechanical life support due to inability to breath effectively to maintain normal levels of oxygen and carbon dioxide in the body. | E0750 | 10080927 | Dependence on ventilator
| Wheezing | A symptom and finding during physical examination, characterized by high-pitched, whistling sound during breathing. it results from narrowing or obstruction of the respiratory airways. | E0751 | 10047924 | Wheezing
| High Oxygen Saturation | A high level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | E2202 | 10033319 | Oxygen saturation high | Investigations and Diagnostic Tests | Respiratory System
| Low Oxygen Saturation | A low level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | E2203 | 10033321 | Oxygen saturation low | Investigations and Diagnostic Tests | Respiratory System
| Pulmonary Embolism | The obstruction of the pulmonary artery or one of its branches by an embolus, sometimes associated with infarction of the lung. | E050303 | 10037377 | Pulmonary embolism | Vascular System | Respiratory System
| Respiratory Distress Syndrome of Newborns | A condition of the newborn marked by dyspnea with cyanosis. | E1516 | 10038692 | Respiratory distress syndrome in newborn | Pregnancy, Childbirth and the Puerperium | Respiratory System
Eye | Capsular Bag Tear | A tear in the capsular bag housing the lens. | E0801 | 10063941 | Posterior capsule tear
| Cataract | Partial or complete opacity of the crystalline lens of one or both eyes that decreases visual acuity and eventually results in blindness. Some cataracts appear in infancy or in childhood, but most develop in older individuals. (Sternberg Diagnostic Surgical Pathology, 3rd ed.) | E0802 | 10007739 | Cataract
| Chemosis | Edema of the mucous membrane of the eyeball and eyelid lining. | E0803 | 10008435 | Chemosis | Eye | Infections
| Conjunctivitis | Inflammation of the conjunctiva of the eye. | E0804 | 10010741 | Conjunctivitis
| Corneal Clouding/Hazing | Clouding of the middle area of the cornea, usually as a result of scarring due to infection, injury, or inflammation. | E0805 | 10052116 | Corneal clouding
| Corneal Decompensation | Inability to maintain corneal integrity; or corneal edema resulting from failure of the corneal endothelium to maintain detumescence. | E0806 | 10052117 | Corneal decompensation
| Corneal Edema | Hazy, swollen cornea. | E0807 | 10011007 | Corneal edema | Eye | Generalized Disorders | Corneal Stromal Edema | Fluid from the aqueous humor enters the corneal stroma and causes swelling. | E080701 | 10054772 | Corneal stromal edema | Eye | Generalized Disorders
[8 empty] | Microcystic Edema | Corneal edema of inflammatory origin associated with contact lens wear and morphologically characterized by presence of epithelial microcysts. | E080702 | 10068529 | Corneal microcystic edema | Eye | Generalized Disorders
| Corneal Epithelial Microcysts | Very small, round vesicles containing fluid and cellular debris observed on the surface of the cornea under slit lamp examination in some types of corneal dystrophy and in wearers of extended wear lenses. | E0808 | 10068144 | Corneal epithelial microcysts
| Corneal Infiltrates | Discrete, small lesions present in the cornea as a result of corneal inflammation and, in some cases, after soft contact lens wear especially extended-wear lenses. | E0809 | 10011022 | Corneal infiltrates
| Corneal Pannus | Superficial vascularization of the cornea with infiltration of granulation tissue. | E0810 | 10033682 | Pannus (corneal)
| Corneal Perforation | A rupture in the corneal tissue due to traumatic or pathologic processes. | E0811 | 10011039 | Corneal perforation | Eye | Procedural Complications
| Corneal Scar | Replacement of corneal tissue with scar tissue as a result of injury to the deeper layers of the cornea. | E0812 | 10011044 | Corneal scar
| Corneal Touch | Complications of intraocular lens surgery; usually refers to intraocular lens contact with cornea, can be intermittent or chronic. | E0813 | 10011055 | Corneal touch
| Corneal Ulcer | Area of epithelial tissue loss from corneal surface; associated with inflammatory cells in the cornea and anterior chamber. | E0814 | 10048492 | Corneal ulcer
| Dry Eye(s) | A syndrome characterized by dryness of the cornea and conjunctiva. It is usually caused by a deficiency in tear production. Symptoms include a feeling of burning eyes and a possible foreign body presence in the eye. | E0815 | 10013774 | Dry eye
| Endophthalmitis | Inflammation of the intraocular fluids or tissues with or without microbial organisms. | E0816 | 10014801 | Endophthalmitis | Eye | Infections
| Excessive Tear Production | Profuse lacrimation. | E0817 | 10015591 | Excess tears
| Eye Infections | An infection of any part of the eye, including the eyelid. | E0818 | 10015929 | Eye infection | Eye | Infections | Intraocular Infection | Infection within the eye. | E081801 | 10054762 | Eye infection intraocular | Eye | Infections
| Eye Injury | Damage to the eye. | E0819 | 10061128 | Eye injury | Eye | Injury | Corneal Abrasion | Injury to the epithelium of the cornea. | E081901 | 10010984 | Corneal abrasion
[8 empty] | Eye Burn | Burn injury, including chemical burn, to any part of the eye. | E081902 | 10015911 | Eye burns | Eye | Injury
[8 empty] | Retinal Detachment | Separation of the inner retina layers from the underlying pigment epithelium. | E081903 | 10038848 | Retinal detachment
[8 empty] | Retinal Injury | Damage to the retina due to toxicity, exposure to bright lights, or trauma. | E081904 | 10057430 | Retinal injury | Eye | Injury
[8 empty] | Retinal Tear | A usually small tearing of the retina occurring when the vitreous separates from the retina. It may lead to retinal detachment. | E081905 | 10038897 | Retinal tear
[8 empty] | Vitreous Detachment | Detachment of the vitreous humor from the retina. | E081906 | 10047650 | Vitreous detachment | Eye | Procedural Complications
| Eye Pain | Painful sensation in the eye. | E0820 | 10015958 | Eye pain | Eye | Generalized Disorders
| Foreign Body Sensation in Eye | Feeling of grittiness or having something in the eye; frequently caused by a foreign body. | E0821 | 10051116 | Foreign body sensation in eyes
| Glaucoma | Increased pressure in the eyeball due to obstruction of the outflow of aqueous humor. | E0822 | 10018304 | Glaucoma
| Halo | Hazy ring around bright lights seen by some patients with refractive error or optical defects, e.g. cataracts, or corneal swelling. | E0823 | 10019099 | Halo vision
| Hyphema | Hemorrhage within the anterior chamber of the eye; bloodshot. | E0824 | 10020924 | Hyphema
| Hypopyon | An accumulation of pus in the anterior chamber of the eye. | E0825 | 10021086 | Hypopyon
| Intraocular Pressure Decreased | Decrease of pressure of the intraocular fluid in the eye. | E0826 | 10022804 | Intraocular pressure decreased
| Intraocular Pressure Increased | Increase of pressure of the intraocular fluid in the eye. | E0827 | 10022806 | Intraocular pressure increased
| Iritis | Inflammation of the iris. | E0828 | 10022955 | Iritis
| Keratitis | Inflammation of the cornea. | E0829 | 10023332 | Keratitis |  |  | Acanthameba Keratitis | Keratitis due to infection by acanthamoba. | E082901 | 10074137 | Acanthameba keratitis
| Macular Edema | Accumulation of intraretinal fluid and protein in the macula, which may result in swelling and decreased central vision. | E0830 | 10054467 | Macular edema | Eye | Generalized Disorders
| Overwear Syndrome | Ocular pain which may be very intense, accompanied by corneal epithelium damage, conjunctival injection, lacrimation, blepharospasm, photophobia and hazy vision following corneal edema caused by overwear of contact lenses, principally the PMMA type. | E0831 | 10080905 | Overwear syndrome
| Pupillary Block | An obstruction of the pupil. | E0832 | 10037520 | Pupillary block
| Red Eye(s) | Lay term applied to any condition with dilation of conjunctival or ciliary blood vessels; innumerable causes. | E0833 | 10038189 | Red eye
| Retinal Degeneration | Degeneration of the retina. | E0834 | 10038845 | Retinal degeneration
| Toxic Anterior Segment Syndrome (TASS) | An acute postoperative inflammatory reaction in which a noninfectious substance enters the anterior segment and induces toxic damage to the intraocular tissues. | E0835 | 10066366 | Toxic anterior segment syndrome
| UGH (Uveitis Glaucoma Hyphema) Syndrome | Clinical triad of uveitis, glaucoma, and hyphema (UGH), especially associated with the insertion of an anterior chamber intraocular lens (IOL). | E0836 | 10068156 | Uveitis-glaucoma-hyphema syndrome
| Uveitis | Inflammation of the uvea. | E0837 | 10046851 | Uveitis
| Visual Disturbances | An interference to normal eyesight without necessarily affecting visual acuity. | E0838 | 10047545 | Visual disturbances |  |  | Flashers | The perception of sudden or brief bursts of light. | E083801 | 10047569 | Visual flashes
[8 empty] | Phosphene Visualization | An objective visual sensation that appears with the eyes closed and in the absence of visual light. | E083802 | 10034936 | Phosphenes | Eye | Generalized Disorders
[8 empty] | Vitreous Floaters | Spots or stringy shapes in vision, often caused by deposits, aggregates, or cell fragments in the vitreous humor or lens. | E083803 | 10047654 | Vitreous floaters
| Visual Impairment | Decreased ability to see to a degree that requires corrective measures | E0839 | 10047571 | Visual impairment |  |  | Blurred Vision | The loss of visual acuity (sharpness of vision). | E083901 | 10005886 | Blurred vision
[8 empty] | Loss of Vision | Severely reduced visual acuity that cannot be corrected, or total loss of vision. | E083902 | 10024881 | Loss of vision
| Vitreous Hemorrhage | Blood extravasation in the vitreous humor. | E0840 | 10047656 | Vitreous hemorrhage
| Vitreous Loss | loss or prolapse of vitreous gel. | E0841 | 10063942 | Vitreous loss
| Vitritis | Inflammatory intraocular reaction with clouding and cells in vitreous; often accompanies inflammation of ciliary body, iris, choroid, or retina. | E0842 | 10047663 | Vitritis
| Zonular Dehiscence | Rupture of the fibrrous strands connecting the ciliary body with the crysaline lens of the eye (usually at the time of cataract surgery). | E0843 | 10068351 | Ciliary zonular dehiscence
Ear and Labyrinth | Acoustic Shock | Neurophysiological and psychological symptoms caused by sudden and unexpected loud noise. | E0901 | 10068150 | Acoustic shock | Ear and Labyrinth | Nervous System
| Acoustic Trauma | Injury to the hearing mechanisms within the inner ear, caused by excessively loud noise. | E0902 | 10000528 | Acoustic trauma
| Hearing Impairment | Partial or complete loss of the ability to detect or understand sounds resulting from damage to the outer, middle, or inner ear structures. | E0903 | 10019245 | Hearing impaired |  |  | Partial Hearing Loss | A condition in which a person partially loses the ability to hear sounds in one or both ears. | E090301 | 10050564 | Partial hearing loss
[8 empty] | Total Hearing Loss | A condition in which a person completely loses the hearing in one or both ears. | E090302 | 10074290 | Complete hearing loss
| Tinnitus | A noise in the ears, such as ringing, buzzing, roaring or clicking in the absence of auditory stimulation. | E0904 | 10043882 | Tinnitus | Ear and Labyrinth | Nervous System
| Vertigo | A feeling of movement, a sensation as if the external world were revolving around the person or as if they were revolving in space (subjective vertigo). Vertigo is medically distinct from dizziness, lightheadedness, and unsteadiness. | E0905 | 10047340 | Vertigo
| Neonatal Hearing Impairment | An abnormality that affects hearing within the first month after birth. | E1512 | 10080921 | Neonatal hearing impairment | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth | Complete Neonatal Hearing Loss | Complete loss of the ability to detect or understand sounds present in an infant within its first month after birth. | E151201 | 10082499 | Complete neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth
[8 empty] | Partial Neonatal Hearing loss | Partial loss of the ability to detect or understand sounds present in an infant within its first month after birth. | E151202 | 10082496 | Partial neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth
Gastrointestinal System | Abdominal Distention | The condition of the abdomen being enlarged or swollen from internal pressure. | E1001 | 10000060 | Abdominal distension
| Abdominal Pain | Painful sensation in the abdominal region. | E1002 | 10000081 | Abdominal pain | Gastrointestinal System | Generalized Disorders
| Achalasia | Failure of the smooth muscle fibers of the gastrointestinal tract to relax at any one point of junction of one part with another. | E1003 | 10000420 | Achalasia
| Ascites | Accumulation of fluid in the peritoneal cavity. The fluid may be serous, hemorrhagic, or the result of tumor metastasis to the peritoneum. | E1004 | 10003445 | Ascites
| Bowel Burn | A thermal injury to the bowels. | E1005 | 10082501 | Intestinal burn | Gastrointestinal System | Injury
| Bowel Perforation | A rupture in the wall of the small or large intestine due to traumatic or pathologic processes. | E1006 | 10006056 | Bowel perforation | Gastrointestinal System | Procedural Complications
| Constipation | Irregular and infrequent or difficult evacuation of the bowels. | E1007 | 10010774 | Constipation
| Diarrhea | Frequent, watery bowel movements. | E1008 | 10012727 | Diarrhea
| Dysphagia/ Odynophagia | A symptom referring to difficult and or painful swallowing. | E1009 | 10013950 | Dysphagia
| Eructate | To belch, the casting of upwind from the stomach. | E1010 | 10015137 | Eructation
| Flatus | Gas or air in the gastrointestinal tract. | E1011 | 10016769 | Flatus
| Gastritis | Inflammation of the stomach. | E1012 | 10017853 | Gastritis
| Gastroesophageal Burn | Thermal, chemical or radiation injury to stomach or esophagus. | E1013 | 10082500 | Gastroesophageal burn | Gastrointestinal System | Injury
| Gastrointestinal Hemorrhage | Bleeding originating from any part of the gastrointestinal tract. | E1014 | 10017960 | Gastrointestinal hemorrhage
| Gastrointestinal Regurgitation | Flow in the opposite direction from normal, as the casting up of undigested food or gas from the stomach, or the backward flowing of blood into the heart, or between the chambers of the heart when a valve is incompetent. | E1015 | 10067171 | Regurgitation
| Hematemesis | Vomiting blood. | E1016 | 10019418 | Hematemesis
| Intra-Abdominal Hemorrhage | Bleeding within the abdominal cavity. | E1017 | 10055291 | Intra-abdominal hemorrhage
| Laceration(s) of Esophagus | A torn, ragged, mangled wound, or an accidental cut of esophagus. | E1018 | 10068164 | Esophageal laceration
| Melena | Abnormally dark tarry feces containing blood, usually from gastrointestinal bleeding. | E1019 | 10027162 | Melena
| Nausea | Upper abdominal discomfort associated with an urge to vomit. | E1020 | 10028813 | Nausea
| Pancreatitis | Inflammation of the pancreas. | E1021 | 10033645 | Pancreatitis | Gastrointestinal System | Infections
| Perforation of Esophagus | The presence of a hole or other type of opening in the esophageal wall through which the contents of the esophagus can pass into the mediastinum. | E1022 | 10034419 | Perforation of esophagus | Gastrointestinal System | Procedural Complications
| Peritoneal Laceration(s) | A cut or tear of the peritoneum. | E1023 | 10068140 | Peritoneal laceration | Gastrointestinal System | Injury
| Peritonitis | Inflammation of the peritoneum due to infection by bacteria or fungi. | E1024 | 10034674 | Peritonitis | Gastrointestinal System | Infections
| Pyrosis/Heartburn | Abdominal discomfort with retrosternal pain usually associated with gastroesophageal reflux. | E1025 | 10019326 | Heartburn
| Rectal Anastomotic Leakage | Leakage due to breakdown of a rectal anastomosis. | E1026 | 10065894 | Rectal anastomotic leak | Gastrointestinal System | Procedural Complications
| Retroperitoneal Hemorrhage | Bleeding originating in the retroperitoneal area. | E1027 | 10038981 | Retroperitoneal hemorrhage
| Salivary Gland Problem | A non-neoplastic or neoplastic (benign or malignant) disorder involving a salivary gland. | E1028 | 10061935 | Salivary gland disorder |  |  | Dry Mouth | The condition of not having enough saliva to keep the mouth moist. | E102801 | 10013781 | Dry mouth
[8 empty] | Salivary Hypersecretion | Excessive saliva production. | E102802 | 10039424 | Salivary hypersecretion
| Stenosis of the esophagus | Narrowing or tightening of the esophagus | E1029 | 10041978 | Stenosis esophageal
| Stomach Ulceration | The formation of sores or a break in the wall of the stomach. | E1030 | 10042116 | Stomach ulcer
| Stomatitis | Inflammation of the oral mucosa. | E1031 | 10042128 | Stomatitis |  |  | Glossitis | Inflammation of the tongue. | E103101 | 10018386 | Glossitis
| Vomiting | Forceful ejection of the contents of the stomach through the mouth. | E1032 | 10047700 | Vomiting
| Difficulty Chewing | Problem with mastication due to pain, discomfort or loss of mobility in the jaw. | E1607 | 10067464 | Chewing difficulty | Musculoskeletal System | Gastrointestinal System
| Increased Intra-Peritoneal Volume (IIPV) | Abnormal build up of fluid in the abdomen. | E1303 | 10082492 | Increased intraperitoneal volume | Kidney and Urinary Tract | Gastrointestinal System
Hepatic and Biliary System | Biliary Cirrhosis | Cirrhosis of the liver caused either by destruction of the intrahepatic bile ducts (primary biliary cirrhosis) or blockage of the extrahepatic bile ducts (secondary biliary cirrhosis). | E1101 | 10004659 | Biliary cirrhosis
| Hepatitis | Inflammation of the liver; usually from a viral infection, but sometimes from toxic agents. | E1102 | 10056522 | Hepatic infection | Hepatic and Biliary System | Infections
| Hyperbilirubinemia | Abnormally high level of bilirubin in the blood. Excess bilirubin is associated with jaundice. | E1103 | 10020582 | Hyperbilirubinemia
| Liver Damage/Dysfunction | Non traumatic damage to the liver parenchyma resulting in hepatic insufficiency. | E1104 | 10019831 | Hepatocellular damage
| Liver Failure | Sustained, and severe damage to the liver with significant loss of function. | E1105 | 10024678 | Liver failure
| Liver Laceration(s) | A cut or tear in the liver. | E1106 | 10068143 | Liver laceration
Metabolism and Nutrition | Dehydration | A condition resulting from the excessive loss of fluid from the body. | E1201 | 10012174 | Dehydration
| Electrolyte Imbalance | Higher or lower than normal values for the serum electrolytes; usually affecting Na, K, Cl, Ca, CO2, BUN. | E1202 | 10014418 | Electrolyte imbalance | Metabolism, and Nutrition | Blood and Lymphatic System | Hypernatremia | Higher than normal levels of sodium in the circulating blood. | E120201 | 10020680 | Hypernatremia | Metabolism, and Nutrition | Blood and Lymphatic System
[8 empty] | Hyponatremia | Lower than normal levels of sodium in the circulating blood. | E120202 | 10021038 | Hyponatremia | Metabolism, and Nutrition | Blood and Lymphatic System
| Feeding Problem | Interruption or disruption of the intake of food/nutrition. | E1203 | 10061148 | Feeding disorder
| Hot Flashes/Flushes | A temporary and/or sudden feeling of intense body warmth, flushing, sometimes accompanied by sweating. | E1204 | 10020407 | Hot flashes
| Hyperglycemia | Abnormally high level of glucose in the blood. | E1205 | 10020639 | Hyperglycemia | Metabolism, and Nutrition | Investigations and Diagnostic Tests | Diabetic Ketoacidosis | A type of metabolic acidosis produced by accumulation of ketone bodies resulting from uncontrolled diabetes mellitus. | E120501 | 10012671 | Diabetic ketoacidosis
| Hypoglycemia | Abnormally low level of glucose in the blood. | E1206 | 10021005 | Hypoglycemia | Metabolism, and Nutrition | Investigations and Diagnostic Tests
| Polydipsia | Extreme thirst accompanied by chronic excessive intake of water. | E1207 | 10036067 | Polydipsia
| Weight Changes | Repetitive loss and regain of body weight. | E1208 | 10049040 | Weight fluctuation
| Decreased Appetite | Lack of desire for food. | E2306 | 10061428 | Decreased appetite | Generalized Disorders | Metabolism, and Nutrition
| Increased Appetite | Intense desire for food. | E2307 | 10021654 | Increased appetite | Generalized Disorders | Metabolism, and Nutrition
| Respiratory Acidosis | A condition in which the blood pH is less than normal, secondary to impaired gas exchange. | E0739 | 10038661 | Respiratory acidosis | Respiratory System | Metabolism, and Nutrition
| Respiratory Alkalosis | A condition in which the blood pH is greater than normal, secondary to impaired gas exchange. | E0740 | 10038664 | Respiratory alkalosis | Respiratory System | Metabolism, and Nutrition
| Uremia | Abnormally high concentration of non-protein nitrogen in the blood. | E1306 | 10046369 | Uremia | Kidney and Urinary Tract | Metabolism, and Nutrition
Kidney and Urinary Tract | Dysuria | Difficult or painful urination. | E1301 | 10013990 | Dysuria
| Hematuria | Blood in the urine. | E1302 | 10019450 | Hematuria
| Increased Intra-Peritoneal Volume (IIPV) | Abnormal build up of fluid in the abdomen. | E1303 | 10082492 | Increased intraperitoneal volume | Kidney and Urinary Tract | Gastrointestinal System
| Micturition Urgency | A sudden compelling urge to urinate. | E1304 | 10027566 | Micturition urgency
| Renal Impairment | An acute or chronic condition where the kidnies are not functioning normally. | E1305 | 10062237 | Renal impairment |  |  | Renal Failure | An acute or chronic condition that is characterized by the inability of the kidneys to adequately filter the blood resulting in uremia and electrolyte imbalances. | E130501 | 10038435 | Renal failure
| Uremia | Abnormally high concentration of non-protein nitrogen in the blood. | E1306 | 10046369 | Uremia | Kidney and Urinary Tract | Metabolism, and Nutrition
| Urethral Stenosis/Stricture | Narrowing of any part of the urethra. With or without affecting the flow of urine. | E1307 | 10065584 | Urethral stenosis
| Urinary Frequency | Urination at short intervals; it may result from increased urine formation, decreased bladder capacity, or lower urinary tract irritation. | E1308 | 10046539 | Urinary frequency
| Urinary Retention | Accumulation of urine within the bladder because of the inability to urinate. | E1309 | 10046555 | Urinary retention
| Urinary Tract Infection | An inflammatory process affecting any part of the urinary tract, most commonly the bladder or the urethra. | E1310 | 10046571 | Urinary tract infection | Kidney and Urinary Tract | Infections | Kidney Infection | An inflammatory process affecting the kidney. | E131001 | 10023424 | Kidney infection | Kidney and Urinary Tract | Infections
Reproductive System and Breast | Abnormal Vaginal Discharge | Secretions from the vagina with a smell, colour, or texture which is different from normal. | E1401 | 10057094 | Vaginal discharge abnormality
| Breast Discomfort/Pain | Having pain, tenderness or discomfort in the breast. | E1402 | 10006298 | Breast pain
| Breast Mass | Abnormality in the breast parenchyma, identified either by palpation or imaging. These abnormalities include calcifications, breast tissue distortion, presence of areas of density, and presence of benign or malignant tumors. | E1403 | 10006272 | Breast mass
| Cervical Changes | Changes such as abnormal, pathological, benign, malignant, etc. | E1404 | 10008351 | Cervix disorder | Reproductive System and Breast | Investigations and Diagnostic Tests
| Dyspareunia | Painful or difficult coitus. | E1405 | 10013941 | Dyspareunia
| Fallopian Tube Perforation | A rupture in the fallopian tube due to traumatic or pathologic processes. | E1406 | 10065790 | Fallopian tube perforation | Reproductive System and Breast | Procedural Complications
| Genital Bleeding | Abnormal bleeding from the genitalia. | E1407 | 10071812 | Genital bleeding |  |  | Intermenstrual Bleeding | Uterine bleeding that occurs during a time not associated with the normal menstrual period. | E140701 | 10022559 | Intermenstrual bleeding
[8 empty] | Menorrhagia | Heavy or large blood loss during menses. | E140702 | 10027313 | Menorrhagia
| Menstrual Irregularities | Deviations from the normal process; e.g. delayed, difficult, profuse, scanty, unusual bleeding, etc. | E1408 | 10027333 | Menstrual irregularity
| Nipple Sensation Changes | Excessive sensitivity or changes in sensitivty in the nipple(s) | E1409 | 10082498 | Nipple sensation changes
| Nipple Ulceration | Ulcer that develops in the nipple. | E1410 | 10029424 | Nipple ulceration
| Pelvic Inflammatory Disease | Any pelvic infection involving the upper female genital tract beyond the cervix. | E1411 | 10034254 | Pelvic inflammatory disease
| Rectovaginal Fistula | Abnormal connetion between rectum and vagina. | E1412 | 10051097 | Rectovaginal fistula
| Sexual Dysfunction | Difficulty, lack of desire, or pain associated with coitus. | E1413 | 10040477 | Sexual dysfunction |  |  | Erectile Dysfunction | Persistent or recurrent inability to achieve or to maintain an erection during sexual activity. | E141301 | 10061461 | Erectile dysfunction
| Toxic Shock Syndrome | A rare acute life-threatening systemic bacterial noncontagious illness caused by staphylococcal exotoxins. | E1414 | 10044248 | Toxic shock syndrome | Reproductive System and Breast | Infections
| Uterine Perforation | A rupture in the uterus due to traumatic or pathologic processes. | E1415 | 10046810 | Uterine perforation | Reproductive System and Breast | Procedural Complications
| Vesicovaginal Fistula | Abnormal connection between vagina and bladder. | E1416 | 10069086 | Vesicovaginal fistula
| Ectopic Pregnancy | Pregnancy condition that happens when there is a developing fetus or embryo outside the uterus. | E1506 | 10014166 | Ectopic pregnancy | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast
| Incompetent Cervix | Cervix that is abnormally prone to dilate without contractions in the second trimester of pregnancy, resulting in premature expulsion of the fetus. | E1509 | 10021632 | Incompetent cervix | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast
| Vaginal Mucosa Damage | Injury to vaginal mucosa. | E1519 | 10080926 | Vaginal mucosal damage | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast
Pregnancy, Childbirth and the Puerperium | Abortion | The unintentional or intentional loss of a pregnancy before 22 weeks gestation. | E1501 | 10000210 | Abortion |  |  | Complete Induced Abortion | Expulsion of all the products of conception following medical or surgical pregnancy termination. | E150101 | 10060928 | Abortion induced complete
[8 empty] | Incomplete Induced Abortion | The uterus is not entirely emptied of products of conception following medical or surgical pregnancy termination. | E150102 | 10053984 | Abortion induced incomplete
[8 empty] | Induced Abortion | Abortion brought on intentionally. | E150103 | 10021716 | Induced abortion
[8 empty] | Missed Abortion | Retention in uterus of an abortus. | E150104 | 10027704 | Missed abortion
[8 empty] | Spontaneous Abortion | Spontaneous expulsion of a human fetus before it is viable. | E150105 | 10041687 | Spontaneous abortion
| Amnionitis | Inflammation of the amnion. | E1502 | 10051242 | Amnionitis
| Cephalohematoma | A subperiosteal hemorrhage limited to the surface of one cranial bone, a usually benign condition seen in the newborn as a result of bone trauma. | E1503 | 10060531 | Cephalohematoma
| Chorioamnionitis | An infection involving chorion, amnion and amniotic fluid; generally, there is also involve placental villi and from decidua. | E1504 | 10008755 | Chorioamnionitis | Pregnancy, Childbirth and the Puerperium | Infections
| Congenital Defect/Deformity | Any abnormality either hereditary or due an influence that happen during the pregnancy until the moment of birth. | E1505 | 10048305 | Birth defects | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System
| Ectopic Pregnancy | Pregnancy condition that happens when there is a developing fetus or embryo outside the uterus. | E1506 | 10014166 | Ectopic pregnancy | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast
| Fetal Distress | Signs or symptoms which may be indicative of distress to a fetus, which may include repetitive variable decelerations, fetal tachycardia or bradycardia, late decelerations, or low biophysical profile. | E1507 | 10016486 | Fetal distress | Pregnancy, Childbirth and the Puerperium | Heart
| Hypoxia in Utero | Caused by conditions such as inadequate placental function (often abruptio placentae), preeclamptic toxicity, prolapse of the umbilical cord, or complications from anesthetic administration. | E1508 | 10022748 | Intra-uterine hypoxia
| Incompetent Cervix | Cervix that is abnormally prone to dilate without contractions in the second trimester of pregnancy, resulting in premature expulsion of the fetus. | E1509 | 10021632 | Incompetent cervix | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast
| Low Apgar Score | Any Apgar score less than seven. This may be the initial or a subsequent apgar score. | E1510 | 10002944 | Apgar score low | Pregnancy, Childbirth and the Puerperium | Investigations and Diagnostic Tests
| Neonatal Deformities | Anatomic abnormality manifested during the neonatal period of physical or mechanical origin. | E1511 | 10082494 | Neonatal deformity | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System
| Neonatal Hearing Impairment | An abnormality that affects hearing within the first month after birth. | E1512 | 10080921 | Neonatal hearing impairment | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth | Complete Neonatal Hearing Loss | Complete loss of the ability to detect or understand sounds present in an infant within its first month after birth. | E151201 | 10082499 | Complete neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth
[8 empty] | Partial Neonatal Hearing loss | Partial loss of the ability to detect or understand sounds present in an infant within its first month after birth. | E151202 | 10082496 | Partial neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth
| Perinatal Brain Injury | Injury to the brain of a newborn infant occurring around the time of birth. | E1513 | 10056392 | Perinatal brain damage | Pregnancy, Childbirth and the Puerperium | Nervous System
| Pregnancy with a Contraceptive Device | The woman became pregnant despite using a contraceptive device. | E1514 | 10063130 | Pregnancy with contraceptive device
| Premature Labor | Delivery of a viable infant before the normal end of gestation. | E1515 | 10036599 | Premature labor
| Respiratory Distress Syndrome of Newborns | A condition of the newborn marked by dyspnea with cyanosis. | E1516 | 10038692 | Respiratory distress syndrome in newborn | Pregnancy, Childbirth and the Puerperium | Respiratory System
| Teratogenic Effects | Appearance of anomalies and/or developmental defects due to exposure to an agent that can disturb the development of an embryo or fetus. | E1517 | 10043275 | Teratogenicity
| Toxins In Children | Children affected by toxins may be due to silicone breast implants and their effects on unborn children and from breastfeeding. | E1518 | 10053487 | Exposure to toxic agent
| Vaginal Mucosa Damage | Injury to vaginal mucosa. | E1519 | 10080926 | Vaginal mucosal damage | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast
| Toxemia | The condition resulting from the spread of bacterial products (toxins) by the bloodstream. | E190103 | 10059857 | Toxemia | Infections | Pregnancy, Childbirth and the Puerperium
Musculoskeletal System | Arthralgia | Pain in a joint. | E1601 | 10003239 | Arthralgia
| Arthritis | An inflammatory process affecting a joint. | E1602 | 10003246 | Arthritis |  |  | Rheumatoid Arthritis | A chronic systemic disease, primarily of the joints, marked by inflammatory changes in the synovial membranes and articular structures, widespread fibrinoid degeneration of the collagen fibers in mesenchymal tissues, and by atrophy and rarefaction of bony structures. | E160201 | 10039073 | Rheumatoid arthritis
| Bone Fracture(s) | A traumatic or pathological injury to the bone in which the continuity of the bone is broken. | E1603 | 10017076 | Fracture | Musculoskeletal System | Injury | Hip Fracture | Traumatic or pathological injury to the hip in which the continuity of either the femoral head, femoral neck, intertrochanteric or subtrochanteric regions is broken. | E160301 | 10020100 | Hip fracture | Musculoskeletal System | Injury
[8 empty] | Limb Fracture | A traumatic break in one or more of the bones in a limb. | E160302 | 10074551 | Limb fracture | Musculoskeletal System | Injury
[8 empty] | Multiple Fractures | Two or more simultaneous traumatic breaks in one or more bones. | E160303 | 10028200 | Multiple fractures | Musculoskeletal System | Injury
[8 empty] | Skull Fracture | A traumatic or pathologic injury to the bones of the skull in which the continuity of the bones of the skull is broken. | E160304 | 10061365 | Skull fracture | Musculoskeletal System | Injury
[8 empty] | Vertebral Fracture | Traumatic or pathological injury to the spine in which the continuity of a vertebral bone is broken. | E160305 | 10047333 | Vertebral fracture | Musculoskeletal System | Injury
| Connective Tissue Disease | A non-neoplastic or neoplastic disorder that affects the connective tissue. | E1604 | 10061087 | Connective tissue disorder
| Cramp(s) /Muscle Spasm(s) | A sustained, sudden and involuntary contraction of a muscle or group of muscles. | E1605 | 10011285 | Cramp |  |  | Abdominal Cramps | An involuntary muscular contraction involving a muscle of the abdomen or of a hollow organ within the abdomen. | E160501 | 10000057 | Abdominal cramps
| Damage to Ligament(s) | An injury to any of the fibrous bands of tissue that comprise an articulation (joint). | E1606 | 10061223 | Ligament injury | Musculoskeletal System | Injury
| Difficulty Chewing | Problem with mastication due to pain, discomfort or loss of mobility in the jaw. | E1607 | 10067464 | Chewing difficulty | Musculoskeletal System | Gastrointestinal System
| Fasciitis | Inflammation of fascia. | E1608 | 10016228 | Fasciitis
| Fibromyositis | Inflammation and fibrous degeneration of a muscle. | E1609 | 10063041 | Fibromyositis | Musculoskeletal System | Infections
| Head Injury | A traumatic injury to the head. | E1610 | 10019196 | Head injury | Musculoskeletal System | Injury
| Hyperextension | Forceful extension of a joint beyond its normal limits. | E1611 | 10062310 | Joint hyperextension
| Inadequate Osseointegration | Insufficient structural and functional connection between living bone and the surface of a synthetic implant. | E1612 | 10068157 | Inadequate osteointegration
| Intervertebral Disc Compression or Protrusion | Degenerative, pathological or traumatic changes in the cartilaginous tissue of the spine leading to unwanted compression or protrusion of the disc. | E1613 | 10055039 | Intervertebral disc compression
| Joint Dislocation | An abnormal separation where two or more bones meet. | E1614 | 10023204 | Joint dislocation | Musculoskeletal System | Injury | Subluxation | Partial dislocation of the bones in a joint. | E161401 | 10042402 | Subluxation | Musculoskeletal System | Injury
| Joint Laxity | Looseness or instability of a joint. | E1615 | 10070874 | Joint laxity
| Loss of Range of Motion | A finding that the flexibility of a joint is below the expected range of normal for that individual. | E1616 | 10048706 | Joint range of motion decreased |  |  | Ankylosis | Immobility of the joint due to the formation of bone, cartilaginous or fibrous tissues around the joints. | E161601 | 10002559 | Ankylosis
[8 empty] | Joint Contracture | Permanent restriction of a joint, usually resulting from prolonged spasticity in a muscle area. | E161602 | 10023201 | Joint contracture
[8 empty] | Muscular Rigidity | An involuntary, persistent state of firm, tense muscles with marked resistance to passive movement. | E161603 | 10028330 | Muscle rigidity
| Malunion of Bone | Faulty healing of bone, resulting in improper anatomical alignment. | E1617 | 10026729 | Malunion of fracture
| Metal Related Pathology | Aseptic fibrosis or local necrosis secondary to metal corrosion and release of wear debris. It's also known as Metallosis. | E1618 | 10081986 | Periprosthetic metallosis
| Muscle Hypotonia | Abnormally low level of muscle tone. | E1619 | 10081180 | Muscle hypotonia
| Muscle/Tendon Damage | Any injury to muscle or tendon tissue including strains. | E1620 | 10028314 | Muscle injury
| Muscle Weakness | A reduction in the strength of one or more muscles. | E1621 | 10028350 | Muscle weakness | Musculoskeletal System | Nervous System
| Myalgia | Pain sensation originating from a muscle or group of muscles. | E1622 | 10028411 | Myalgia | Musculoskeletal System | Generalized Disorders
| Neck Pain | Pain in the neck area. | E1623 | 10028836 | Neck pain | Musculoskeletal System | Generalized Disorders
| Neck Stiffness | Limited mobility of the neck, usually accompanied by pain. | E1624 | 10028843 | Neck stiffness
| Non-union Bone Fracture | Broken bone that doesn't heal. | E1625 | 10029746 | Nonunion of fracture
| Ossification | The formation of bone or of a bony substance; the conversion of fibrous tissue or of cartilage into bone or a bony substance. | E1626 | 10065453 | Ectopic ossification
| Osteolysis | Dissolution of bone; applied especially to the removal or loss of the calcium of bone. | E1627 | 10031248 | Osteolysis
| Osteomyelitis | Infection of bone or bone marrow. | E1628 | 10031252 | Osteomyelitis | Musculoskeletal System | Infections
| Osteopenia/ Osteoporosis | Decreased calcification or density of bone tissue. | E1629 | 10049088 | Osteopenia
| Polymyositis | An idiopathic inflammatory disorder affecting the muscles. It presents with symmetrical proximal muscle weakness and elevated skeletal muscle enzymes. | E1630 | 10036102 | Polymyositis
| Spinal Column Injury | Damage to the bones of the spine. | E1631 | 10069689 | Spinal column injury | Musculoskeletal System | Injury
| Synovitis | Inflammation of a synovial membrane. | E1632 | 10042868 | Synovitis | Musculoskeletal System | Infections
| Tooth Fracture | Breakage of any portion of the tooth. | E1633 | 10062544 | Tooth fracture | Musculoskeletal System | Injury
| Unequal Limb Length | Discrepancy between the lengths of the lower or upper extremities. | E1634 | 10065738 | Unequal limb length
| Ambulation Difficulties | Any persistent problem with walking or gait. | E2302 | 10047810 | Walking difficulty | Generalized Disorders | Musculoskeletal System
| Congenital Defect/Deformity | Any abnormality either hereditary or due an influence that happen during the pregnancy until the moment of birth. | E1505 | 10048305 | Birth defects | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System
| Neonatal Deformities | Anatomic abnormality manifested during the neonatal period of physical or mechanical origin. | E1511 | 10082494 | Neonatal deformity | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System
Skin and Subcutaneous Tissue | Abrasion | Superficial damage to the skin caused by rubbing or scraping. | E1701 | 10080919 | Abrasion | Skin and Subcutaneous Tissue | Injury
| Angioedema | Rapid swelling of the deep layers of the skin due to transient vascular leakage of serous fluid. | E1702 | 10002424 | Angioedema | Skin and Subcutaneous Tissue | Generalized Disorders
| Blister | A collection of fluid underneath the top layer of skin. | E1703 | 10005191 | Blister
| Burn(s) | An injury or tissue damage caused by exposure to caustic chemicals, extreme heat, extreme cold or excessive radiation. | E1704 | 10006634 | Burn | Skin and Subcutaneous Tissue | Injury | Caustic/Chemical Burns | A burn caused by exposure to chemical substances. | E170401 | 10008420 | Chemical burn | Skin and Subcutaneous Tissue | Injury
[8 empty] | Cryogenic Burn | Local tissue damage due to exposure to extreme cold. | E170402 | 10081985 | Cold burn | Skin and Subcutaneous Tissue | Injury
[8 empty] | Radiation Burn | A burn caused by exposure to x-ray, radium, sunlight, ionizing radiation or any other type of radiant energy. | E170403 | 10063640 | Radiation burn | Skin and Subcutaneous Tissue | Injury
[8 empty] | Superficial (First Degree) Burn | A burn that affects the epidermis only, causing erythema without blistering. | E170404 | 10016700 | First degree burns | Skin and Subcutaneous Tissue | Injury
[8 empty] | Partial thickness (Second Degree) Burn | A burn that affects the epidermis and the dermis. | E170405 | 10039798 | Second degree burns | Skin and Subcutaneous Tissue | Injury
[8 empty] | Full thickness (Third Degree) Burn | A burn to the skin that extends into deeper layers and underlying tissues. | E170406 | 10043441 | Third degree burns | Skin and Subcutaneous Tissue | Injury
| Burning Sensation | A sensation of stinging or heat, not necessarily accompanied by redness or physical signs of irritation. | E1705 | 10006784 | Burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders | Skin Burning Sensation | A sensation of dermal stinging or heat, not necessarily accompanied by redness or physical signs of skin irritation. | E170501 | 10054786 | Skin burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders
| Cancer Cells Dissemination | Spreading benign or malignant tissue during medical procedures. | E1706 | 10082075 | Iatrogenic metastasis
| Impaired Healing | Improper body tissue repair due to factors affecting one or more of the phases of wound healing. | E1707 | 10021519 | Impaired healing
| Itching Sensation | An irritating feeling that produces the desire to scratch. | E1708 | 10023084 | Itching
| Jaundice | Yellow pigmentation of the skin, mucous membranes, and the eyes due to hyperbilirubinemia. | E1709 | 10023126 | Jaundice
| Localized Skin Lesion | A pathologic process that affects the skin and is confined to a specific area. | E1710 | 10040882 | Skin lesion
| Peeling | A peeling off or loss of epidermis, as in sunburn, postscarlatinal peeling, or toxic epidermal necrolysis. | E1711 | 10034218 | Peeling
| Phototoxicity | A nonimmunologic, chemically induced type of photosensitivity. | E1712 | 10034977 | Phototoxicity
| Pocket Erosion | Erosion of patient's tissue pocket which houses a device. | E1713 | 10035771 | Pocket erosion | Skin and Subcutaneous Tissue | Injury
| Rash | An eruption in the skin which effects its appearance and/or texture. | E1714 | 10037844 | Rash
| Scar Tissue | Formation of new tissue formed in the healing of a wound. | E1715 | 10039580 | Scar
| Skin Discoloration | any abnormal change in skin coloration. | E1716 | 10040828 | Skin discoloration |  |  | Erythema | Red discoloration of the skin. | E171601 | 10015150 | Erythema
| Skin Disorders | Affecting the skin. | E1717 | 10040831 | Skin disorder
| Skin Erosion | A gradual breakdown or very shallow ulceration of the skin which involves only the epidermis and heals without scarring. | E1718 | 10040840 | Skin erosion | Skin and Subcutaneous Tissue | Injury
| Skin Infection | An inflammatory process affecting the skin, caused by bacteria, viruses, parasites, or fungi. | E1719 | 10040872 | Skin infection | Skin and Subcutaneous Tissue | Infections
| Skin Inflammation/ Irritation | An inflammatory process affecting the skin. Signs include red rash, itching, and blister formation. Representative examples are contact dermatitis, atopic dermatitis, and seborrheic dermatitis. | E1720 | 10062249 | Skin inflammation |  |  | Abscess | An inflammatory process characterized by the accumulation of pus within a newly formed tissue cavity. | E172001 | 10000269 | Abscess | Skin and Subcutaneous Tissue | Infections
[8 empty] | Cellulitis | Inflammation of the dermis and subcutaneous tissues caused by a bacterial infection. | E172002 | 10007882 | Cellulitis | Skin and Subcutaneous Tissue | Infections
[8 empty] | Contact Dermatitis | An inflammatory skin condition caused by direct contact between the skin and either an irritating substance or an allergen. | E172003 | 10056540 | Dermatitis irritant contact
[8 empty] | Eczema | A form of dermatitis characterized by red, itchy, scaly, or crusty patches that can be chronic or intermittent. | E172004 | 10014184 | Eczema
[8 empty] | Radiodermatitis | A cutaneous inflammatory reaction occurring as a result of exposure to biologically effective levels of ionizing radiation. | E172005 | 10061103 | Dermatitis radiation
[8 empty] | Urticaria | A transient, itchy skin eruption characterized by wheals with pale interiors and red margins. | E172006 | 10046735 | Urticaria
[8 empty] | Wheal(s) | A smooth, slightly elevated area on the body surface, which is redder or paler than the surrounding skin. | E172007 | 10047918 | Wheals
[8 empty] | Granuloma | An inflammatory reaction usually caused by infectious organisms, foreign bodies, or cholesterol deposits. It is characterized by the presence of epithelioid histiocytes and chronic inflammation. | E2317 | 10018691 | Granuloma | Generalized Disorders | Skin and Subcutaneous Tissue
| Skin Tears | A break in skin tissue caused by friction or shearing. | E1721 | 10048704 | Skin tear
| Subcutaneous Nodule | A small palpable abnormal mass in the subcutaneous tissues. | E1722 | 10042348 | Subcutaneous nodule
| Wrinkling | A folding, ridge or creasing of the skin. | E1723 | 10048043 | Wrinkling
| Easy Bruising | Appearance of bruises spontaneously or following minimal contact. | E2005 | 10076258 | Tendency to bruise easily | Injury | Skin and Subcutaneous Tissue
Neoplasms Benign, Malignant and Unspecified | Cancer | A group of diseases in which abnormal cells divide without control, invade nearby tissues and may also spread to other parts of the body through the blood and lymph systems. | E1801 | 10007050 | Cancer |  |  | Breast Cancer | A malignant neoplasm that develops or arises in breast tissue. | E180101 | 10006187 | Breast cancer
[8 empty] | Breast Implant Associated Anaplastic Large Cell Lymphoma (BIA ALCL) | A peripheral (mature) T-cell lymphoma, consisting of usually large anaplastic, CD30 positive cells. The majority of cases are positive for the anaplastic large cell lymphoma (ALK) protein. | E180102 | 10082495 | Breast implant-associated anaplastic large cell lymphoma
[8 empty] | Leiomyosarcoma | A malignant tumor of smooth muscle cells that can arise almost anywhere in the body but is most common in the uterus, abdomen, or pelvis. | E180103 | 10024189 | Leiomyosarcoma
[8 empty] | Lymphoma | A malignant (clonal) proliferation of B- lymphocytes or T- lymphocytes which involves the lymph nodes, bone marrow and/or extranodal sites. This category includes Non-Hodgkin lymphomas and Hodgkin lymphomas. | E180104 | 10025310 | Lymphoma
| Cyst(s) | A sac-like closed membranous structure that may be empty or contain fluid or amorphous material. | E1802 | 10011732 | Cyst
| Nodule | A small lump, swelling or collection of tissue. | E1803 | 10054107 | Nodule
| Solid Tumour | Abnormal mass of tissue that usually does not contain cysts or liquid areas. Solid tumours may be benign (not cancerous), or malignant (cancerous). Different types of solid tumours are named for the type of cells that form them. If tumours are associated with those under the term “cancer”, these terms should be used. | E1804 | 10065252 | Solid tumor
Infections | Bacterial Infection | An acute infectious disorder caused by gram positive or gram negative bacteria. | E1901 | 10060945 | Bacterial infection |  |  | Drug Resistant Bacterial Infection | An infection that is resistant to antibiotics and antimicrobials, for example, MRSA or VRE. | E190101 | 10004052 | Bacterial resistance
[8 empty] | Pyogenic Infection | An infection caused by pus-producing organisms. | E190102 | 10021861 | Infection pyogenic
[8 empty] | Toxemia | The condition resulting from the spread of bacterial products (toxins) by the bloodstream. | E190103 | 10059857 | Toxemia | Infections | Pregnancy, Childbirth and the Puerperium
| Fungal Infection | An infection caused by a fungus. | E1902 | 10017533 | Fungal infection
| Sexually Transmitted Infection | A disease acquired predominantly through sexual contact. | E1903 | 10068720 | Sexually transmitted infection
| Subclinical Infection | Infection associated with no detectable symptoms but caused by microorganisms capable of producing easily recognizable diseases. | E1904 | 10068163 | Subclinical infection
| Transmissible Spongiform Encephalopathy(TSE) | A group of rare degenerative brain disorders associated with orions for example Creutzfeldt-Jakob disease (CJD) or Bovine Spongiform Encephalopathy (BSE). | E1905 | 10080750 | Transmissible spongiform encephalopathy
| Unspecified Infection | A disorder resulting from the presence and activity of a microbial, viral, fungal, or parasitic agent. It can be transmitted by direct or indirect contact. | E1906 | 10021789 | Infection
| Viral Infection | Any disease caused by a virus. | E1907 | 10047461 | Viral infection |  |  | Cytomegalovirus (CMV) Infection | A genus of the family herpesviridae, subfamily betaherpesvirinae, infecting the salivary glands, liver, spleen, lungs, eyes, and other organs, in which they produce characteristically enlarged cells with intranuclear inclusions. | E190701 | 10011831 | Cytomegalovirus infection
| Abscess | An inflammatory process characterized by the accumulation of pus within a newly formed tissue cavity | E172001 | 10000269 | Abscess | Skin and Subcutaneous Tissue | Infections
| Cellulitis | Inflammation of the dermis and subcutaneous tissues caused by a bacterial infection. | E172002 | 10007882 | Cellulitis | Skin and Subcutaneous Tissue | Infections
| Chemosis | Edema of the mucous membrane of the eyeball and eyelid lining. | E0803 | 10008435 | Chemosis | Eye | Infections
| Chorioamnionitis | An infection involving chorion, amnion and amniotic fluid; generally, there is also involve placental villi and from decidua. | E1504 | 10008755 | Chorioamnionitis | Pregnancy, Childbirth and the Puerperium | Infections
| Encephalitis | An acute inflammatory process affecting the brain parenchyma. | E0114 | 10014581 | Encephalitis | Nervous System | Infections
| Endocarditis | Inflammation of the endocardium. | E0610 | 10014665 | Endocarditis | Heart | Infections
| Endophthalmitis | Inflammation of the intraocular fluids or tissues with microbial organisms. | E0816 | 10014801 | Endophthalmitis | Eye | Infections
| Eye Infections | An infection of any part of the eye, including the eyelid | E0818 | 10015929 | Eye infection | Eye | Infections | Intraocular Infection | Infection within the eye. | E081801 | 10054762 | Eye infection intraocular | Eye | Infections
| Fibromyositis | Inflammation and fibrous degeneration of a muscle. | E1609 | 10063041 | Fibromyositis | Musculoskeletal System | Infections
| Hepatitis | Inflammation of the liver; usually from a viral infection, but sometimes from toxic agents. | E1102 | 10056522 | Hepatic infection | Hepatic and Biliary System | Infections
| Meningitis | A disorder characterized by acute inflammation of the meninges of the brain and/or spinal cord. | E0121 | 10027199 | Meningitis | Nervous System | Infections
| Osteomyelitis | Infection of bone or bone marrow. | E1628 | 10031252 | Osteomyelitis | Musculoskeletal System | Infections
| Pancreatitis | Inflammation of the pancreas. | E1021 | 10033645 | Pancreatitis | Gastrointestinal System | Infections
| Peritonitis | Inflammation of the peritoneum due to infection by bacteria or fungi. | E1024 | 10034674 | Peritonitis | Gastrointestinal System | Infections
| Pneumonia | An acute and/or chronic inflammation focally or diffusely affecting the lung parenchyma. | E0733 | 10035664 | Pneumonia | Respiratory System | Infections | Bronchopneumonia | Acute inflammation of the walls of the terminal bronchioles. | E073301 | 10006469 | Bronchopneumonia | Respiratory System | Infections
| Post Operative Wound Infection | Infection of a surgical skin incision. | E2115 | 10036410 | Postoperative wound infection | Procedural Complications | Infections
| Post Traumatic Wound Infection | Infection of a wound following trauma. This does not include post surgical wound infection. | E2012 | 10036317 | Post-traumatic wound infection | Injury | Infections
| Respiratory Tract Infection | An infectious disease affecting the respiratory tract. | E0744 | 10062352 | Respiratory tract infection | Respiratory System | Infections
| Sepsis | The presence of pathogenic microorganisms in the blood stream causing a rapidly progressing systemic reaction that may lead to shock. Symptoms include fever, chills, tachycardia, and increased respiratory rate. It is a medical emergency that requires urgent medical attention. | E0306 | 10040047 | Sepsis | Blood and Lymphatic System | Infections
| Skin Infection | An inflammatory process affecting the skin, caused by bacteria, viruses, parasites, or fungi. | E1719 | 10040872 | Skin infection | Skin and Subcutaneous Tissue | Infections
| Synovitis | Inflammation of a synovial membrane. | E1632 | 10042868 | Synovitis | Musculoskeletal System | Infections
| Toxic Shock Syndrome | A rare acute life-threatening systemic bacterial noncontagious illness caused by staphylococcal exotoxins. | E1414 | 10044248 | Toxic shock syndrome | Respiratory System | Infections
| Urinary Tract Infection | An inflammatory process affecting any part of the urinary tract, most commonly the bladder or the urethra. | E1310 | 10046571 | Urinary tract infection | Kidney and Urinary Tract | Infections | Kidney Infection | An inflammatory process affecting the kidney. | E131001 | 10023424 | Kidney infection | Kidney and Urinary Tract | Infections
Injury | Bone Shedding Debris | Shedding of accumulated bone fragments. | E2001 | 10006011 | Bone shedding debris
| Bruise/Contusion | Injury of the soft tissues without breaking the skin that is characterized by blood leakage into the surrounding area. | E2002 | 10006502 | Bruise |  |  | Myocardial Contusion | A bruise to the heart. | E0615 | 10068147 | Myocardial contusion | Heart | Injury
| Chemical Exposure | Contact with a chemical substance through touch, inhalation, or ingestion. | E2003 | 10068154 | Chemical exposure
| Crushing Injury | Traumatic compression of the body or a part of the body. | E2004 | 10061097 | Crush injury
| Easy Bruising | Appearance of bruises spontaneously or following minimal contact. | E2005 | 10076258 | Tendency to bruise easily | Injury | Skin and Subcutaneous Tissue
| Erosion | An eating away or breakdown of any type of external or internal human tissue including but not limited to skin, teeth, mucosa, or somatic, which involves only the outer tissue layer. When tissue surrounds an implanted device, the tissue breakdown may result in migration and loss of the implant material and may result in further complications such as infection or abscess. | E2006 | 10068527 | Tissue erosion associated with device |  |  | Pocket Erosion | Erosion of patient's tissue pocket which houses a device. | E200601 | 10035771 | Pocket erosion | Skin and Subcutaneous Tissue | Injury
[8 empty] | Skin Erosion | A gradual breakdown or very shallow ulceration of the skin which involves only the epidermis and heals without scarring. | E200602 | 10040840 | Skin erosion | Skin and Subcutaneous Tissue | Injury
| Fall | A sudden movement downward, usually resulting in injury. | E2007 | 10016173 | Fall
| Foreign Body In Patient | An occurrence where any object including device or fragments is left unintentionally in the body. | E2008 | 10070245 | Foreign body | Injury | Procedural Complications | Device Embedded In Tissue or Plaque | Device and/or fragments of device are embedded in patient's tissue and/or plaque. | E200801 | 10074498 | Embedded device | Injury | Procedural Complications
| Laceration(s) | A cut or tear in any tissue. | E2009 | 10023572 | Laceration
| Needle Stick/Puncture | A penetrating stab wound from a needle (or other sharp object) that may result in exposure to blood or other body fluids. | E2010 | 10028895 | Needle stick/puncture
| Physical Entrapment | Patient becomes entangled or trapped in a device. | E2011 | 10082491 | Patient entrapped in device
| Post Traumatic Wound Infection | Infection of a wound following trauma. This does not include post surgical wound infection. | E2012 | 10036317 | Post-traumatic wound infection | Injury | Infections
| Rupture | Forcible tearing or disruption of tissue. | E2013 | 10080972 | Tissue rupture
| Tissue Breakdown | A process caused by injury or disease or aging and characterized by a partial/complete disruption of tissue. The disruption may be temporary or permanent and may be associated with either (1) impairment or (2) loss of function of the affected organ(s) or system(s). | E2014 | 10068521 | Tissue breakdown associated with device |  |  | Pressure Sores | Death of tissue due to external pressure. | E201401 | 10036651 | Pressure sore
| Unspecified Tissue Injury | Injury to healthy tissues. | E2015 | 10080903 | Tissue injury
| Abrasion | Superficial damage to the skin caused by rubbing or scraping. | E1701 | 10080919 | Abrasion | Skin and Subcutaneous Tissue | Injury
| Bone Fracture(s) | A traumatic or pathological injury to the bone in which the continuity of the bone is broken. | E1603 | 10017076 | Fracture | Musculoskeletal System | Injury | Hip Fracture | Traumatic or pathological injury to the hip in which the continuity of either the femoral head, femoral neck, intertrochanteric or subtrochanteric regions is broken. | E160301 | 10020100 | Hip fracture | Musculoskeletal System | Injury
[8 empty] | Limb Fracture | A traumatic break in one or more of the bones in a limb. | E160302 | 10074551 | Limb fracture | Musculoskeletal System | Injury
[8 empty] | Multiple Fractures | Two or more simultaneous traumatic breaks in one or more bones. | E160303 | 10028200 | Multiple fractures | Musculoskeletal System | Injury
[8 empty] | Skull Fracture | A traumatic or pathologic injury to the bones of the skull in which the continuity of the bones of the skull is broken. | E160304 | 10061365 | Skull fracture | Musculoskeletal System | Injury
[8 empty] | Vertebral Fracture | Traumatic or pathological injury to the spine in which the continuity of a vertebral bone is broken. | E160305 | 10047333 | Vertebral fracture | Musculoskeletal System | Injury
| Brain Injury | Damage to the brain. | E0102 | 10060690 | Traumatic brain injury | Nervous System | Injury | Encephalocele | Hernia of brain substance and meninges through a congenital or traumatic opening of the skull. | E010201 | 10014617 | Encephalocele | Nervous System | Injury
| Burn(s) | An injury or tissue damage caused by exposure to caustic chemicals, extreme heat, extreme cold or excessive radiation. | E1704 | 10006634 | Burn | Skin and Subcutaneous Tissue | Injury | Bowel Burn | A thermal injury to the bowels. | E1005 | 10082501 | Intestinal burn | Gastrointestinal System | Injury
[8 empty] | Caustic/Chemical Burns | A burn caused by exposure to chemical substances. | E170401 | 10008420 | Chemical burn | Skin and Subcutaneous Tissue | Injury
[8 empty] | Cryogenic Burn | Local tissue damage due to exposure to extreme cold. | E170402 | 10081985 | Cold burn | Skin and Subcutaneous Tissue | Injury
[8 empty] | Gastroesophageal Burn | Thermal, chemical or radiation injury to stomach or esophagus. | E1013 | 10082500 | Gastroesophageal burn | Gastrointestinal System | Injury
[8 empty] | Radiation Burn | A burn caused by exposure to x-ray, radium, sunlight, ionizing radiation or any other type of radiant energy. | E170403 | 10063640 | Radiation burn | Skin and Subcutaneous Tissue | Injury
[8 empty] | Superficial (First Degree) Burn | A burn that affects the epidermis only, causing erythema without blistering. | E170404 | 10016700 | First degree burns | Skin and Subcutaneous Tissue | Injury
[8 empty] | Partial thickness (Second Degree) Burn | A burn that affects the epidermis and the dermis. | E170405 | 10039798 | Second degree burns | Skin and Subcutaneous Tissue | Injury
[8 empty] | Full thickness (Third Degree) Burn | A burn to the skin that extends into deeper layers and underlying tissues. | E170406 | 10043441 | Third degree burns | Skin and Subcutaneous Tissue | Injury
| Choking | Interference with respritation by ostruction of the larynx or trachea. | E0710 | 10008589 | Choking | Respiratory System | Injury
| Concussion | Traumatic brain injury as a result of the action of a mechanical force on the head. | E0108 | 10010254 | Concussion | Nervous System | Injury
| Damage to Ligament(s) | An injury to any of the fibrous bands of tissue that comprise an articulation (joint). | E1606 | 10061223 | Ligament injury | Musculoskeletal System | Injury
| Electric Shock | A shock caused by a brief, unintentional exposure to an electric current from a device. | E2104 | 10014357 | Electric shock | Procedural Complications | Injury | Shock from Patient Lead(s) | Unintended electric shock from patient leads. This includes an electric shock to a healthcare professional or a patient receiving a shock to an incorrect area of their body | E210401 | 10075322 | Device inappropriate shock delivery | Procedural Complications | Injury
| Exposure to Body Fluids | Exposed to or contacted with blood or body fluids of another person. | E2105 | 10068528 | Exposure to device contaminated with body fluid | Procedural Complications | Injury
| Eye injury | Damage to the eye. | E0819 | 10061128 | Eye injury | Eye | Injury | Eye Burn | Burn injury, including chemical burn, to any part of the eye. | E081902 | 10015911 | Eye burns | Eye | Injury
[8 empty] | Retinal Injury | Damage to the retina due to advanced age, toxicity, exposure to bright lights, or trauma. | E081904 | 10057430 | Retinal injury | Eye | Injury
| Head Injury | A traumatic injury to the head. | E1610 | 10019196 | Head injury | Musculoskeletal System | Injury
| Joint Dislocation | An abnormal separation where two or more bones meet. | E1614 | 10023204 | Joint dislocation | Musculoskeletal System | Injury | Subluxation | Partial dislocation of the bones in a joint. | E161401 | 10042402 | Subluxation | Musculoskeletal System | Injury
| Nerve Damage | Injury to nervous tissue. | E0123 | 10029177 | Nerve damage | Nervous System | Injury | Optical Nerve Damage | Injury to optical nerve. | E012301 | 10030938 | Optic nerve injury | Nervous System | Injury
| Nervous System Injury | Injury to the central or peripheral nervous system. | E0124 | 10080898 | Nervous system injury | Nervous System | Injury | Spinal Cord Injury | Traumatic damage of the spinal cord. | E012401 | 10041552 | Spinal cord injury | Nervous System | Injury
| Peripheral Nervous Injury | Injury to the peripheral nervous system. | E0128 | 10034586 | Peripheral nerve injury | Nervous System | Injury
| Peritoneal Laceration(s) | A cut or tear of the peritoneum. | E1023 | 10068140 | Peritoneal laceration | Gastrointestinal System | Injury
| Radiation Sickness Syndrome | The complex of symptoms characterizing the disease known as radiation injury, resulting from excessive exposure of the whole body (or large part) to ionizing radiation. | E2116 | 10068142 | Radiation sickness syndrome | Procedural Complications | Injury
| Spinal Column Injury | Damage to the bones of the spine. | E1631 | 10069689 | Spinal column injury | Musculoskeletal System | Injury
| Strangulation | Interference with respiration by compression of the larynx, laryngopharynx or trachea,as well as of the carotid arteria or jugular veins. | E0749 | 10042169 | Strangulation | Respiratory System | Injury
| Tooth Fracture | Breakage of any portion of the tooth. | E1633 | 10062544 | Tooth fracture | Musculoskeletal System | Injury
Procedural Complications | Adhesion(s) | A fibrous band of tissue that connects normally separate body regions. | E2101 | 10059837 | Adhesion | Procedural Complications | Generalized Disorders
| Awareness during Anaesthesia | A state of consciousness that permits the formation of explicit and implicit memories while under general anesthesia. | E2102 | 10054330 | Awareness during anesthesia
| Device Overstimulation of Tissue | A level of stimulation that is greater than the intended level. | E2103 | 10068528 | Device overstimulation of tissue
| Electric Shock | A shock caused by a brief, unintentional exposure to an electric current from a device. | E2104 | 10014357 | Electric shock | Procedural Complications | Injury | Shock from Patient Lead(s) | Unintended electric shock from patient leads. This includes an electric shock to a healthcare professional or a patient receiving a shock to an incorrect area of their body. | E210401 | 10075322 | Device inappropriate shock delivery | Procedural Complications | Injury
| Exposure to Body Fluids | Exposed to or contacted with blood or body fluids of another person. | E2105 | 10068528 | Exposure to device contaminated with body fluid | Procedural Complications | Injury
| External Prosthetic Device Pain | Pain that is associated with an external ariticial replacement device. | E2106 | 10059057 | Medical device pain | Procedural Complications | Generalized Disorders
| Failure of Implant | Malfunction of a medical implant. | E2107 | 10016153 | Failure of implant
| Failure to Anastomose | Failure of a surgically-induced connection between tubular structures in the body. | E2108 | 10016164 | Failure to anastomose
| Implant Pain | Pain localized to the site of the implanted device. | E2109 | 10063782 | Implant site pain | Procedural Complications | Generalized Disorders
| Inadequate Pain Relief | Continuing discomfort due to a lack or failure of a medical approach to treat pain. | E2110 | 10079788 | Inadequate pain relief | Procedural Complications | Generalized Disorders
| Injection Site Reaction | An intense reaction (usually immunologic) developing at the site of injection | E2111 | 10022095 | Injection site reaction
| Intraoperative Pain | Pain occurring during a surgical operation. | E2112 | 10064882 | Procedural pain | Procedural Complications | Generalized Disorders
| Oversedation | A level of sedation that is greater than the intended level. | E2113 | 10033305 | Oversedation
| Perforation | A hole or opening made through a membrane or other tissue or material. | E2114 | 10076705 | Perforation |  |  | Internal Organ Perforation | A rupture in the organ(s) due to traumatic or pathologic processes. | E211401 | 10080922 | Perforation of organ
[8 empty] | Bowel Perforation | A rupture in the wall of the small or large intestine due to traumatic or pathologic processes. | E1006 | 10006056 | Bowel perforation | Gastrointestinal System | Procedural Complications
[8 empty] | Cardiac Perforation | The presence of an acquired hole in the heart. | E0604 | 10058039 | Cardiac perforation | Heart | Procedural Complications
[8 empty] | Corneal Perforation | A rupture in the corneal tissue due to traumatic or pathologic processes. | E0811 | 10011039 | Corneal perforation | Eye | Procedural Complications
[8 empty] | Fallopian Tube Perforation | A rupture in the fallopian tube due to traumatic or pathologic processes. | E1406 | 10065790 | Fallopian tube perforation | Reproductive System | Procedural Complications
[8 empty] | Great Vessel Perforation | Damage to the venae cavae, pulmonary artery, pulmonary veins or aorta resulting from a full thickness disruption in the integrity of the vessel wall. | E051101 | 10080896 | Perforation of great vessels | Vascular system | Procedural Complications
[8 empty] | Mitral Perforation | A rupture in the mitral valve tissue due to traumatic or pathologic processes. | E0614 | 10068138 | Mitral perforation | Heart | Procedural Complications
[8 empty] | Perforation of Esophagus | The presence of a hole or other type of opening in the esophageal wall through which the contents of the esophagus can pass into the mediastinum. | E1022 | 10034419 | Perforation of esophagus | Gastrointestinal System | Procedural Complications
[8 empty] | Perforation of Vessels | Damage to a blood vessel resulting from a full thickness disruption in the integrity of the vessel wall. | E0511 | 10068149 | Vessel perforation | Vascular system | Procedural Complications
[8 empty] | Sinus Perforation | A rupture in the sinus tissue due to traumatic or pathologic processes. | E0745 | 10040748 | Sinus perforation | Respiratory System | Procedural Complications
[8 empty] | Uterine Perforation | A rupture in the uterus due to traumatic or pathologic processes. | E1415 | 10046810 | Uterine perforation | Reproductive System | Procedural Complications
| Post Operative Wound Infection | Infection of a surgical skin incision. | E2115 | 10036410 | Postoperative wound infection | Procedural Complications | Infections
| Radiation Sickness Syndrome | The complex of symptoms characterizing the disease known as radiation injury, resulting from excessive exposure of the whole body (or large part) to ionizing radiation. | E2116 | 10068142 | Radiation sickness syndrome | Procedural Complications | Injury
| Twiddlers Syndrome | Dislodgement, breakdown, or other malfunction of an artificial cardiac pacemaker, chemotherapy port, drip infusion valve, or similar implanted diagnostic or therapeutic device as a result of unconscious or habitual manipulation by the patient. | E2117 | 10045184 | Twiddler's syndrome
| Unintended Extubation | The removal of a previously inserted tube from a hollow organ or passageway, often from the airway. | E2118 | 10066823 | Unintended endotracheal extubation
| Unintended Radiation Exposure | The unintentional exposure to radiation of a person, or a part of a persons body. This might include exposure of a healtchare professional, or the exposure of the wrong part of the body of a patient who was intended to receive therapeutic or diagnostic radiation. This includes both ionizing and non ionizing radiation. | E2119 | 10063155 | Radiation exposure
| Foreign Body In Patient | An occurrence where any object including device or fragments is left unintentionally in the body. | E2008 | 10070245 | Foreign body | Injury | Procedural Complications | Device Embedded In Tissue or Plaque | Device and/or fragments of device are embedded in patient's tissue and/or plaque. | E200801 | 10074498 | Embedded device | Injury | Procedural Complications
| Physical Asymmetry | Lack or absence of balanced proportions between parts or organs of the body. | E2332 | - | - | Generalized Disorders | Procedural Complications
| Rectal Anastomotic Leakage | Leakage due to breakdown of a rectal anastomosis. | E1026 | 10065894 | Rectal anastomotic leak | Gastrointestinal System | Procedural Complications
| Vitreous Detachment | Detachment of the vitreous humor from the retina | E081906 | 10047650 | Vitreous detachment | Eye | Procedural Complications
| Wound Dehiscence | A complication in which the wound splits or separates after being surgically closed. | E2340 | 10048031 | Wound dehiscence | Generalized Disorders | Procedural Complications
Investigations and Diagnostic Tests | Abnormal Blood Gases | Abnormal Blood Gas Measurements. | E2201 | 10005539 | Blood gases abnormal
| High Oxygen Saturation | A high level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | E2202 | 10033319 | Oxygen saturation high | Investigations and Diagnostic Tests | Respiratory System
| Low Oxygen Saturation | A low level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | E2203 | 10033321 | Oxygen saturation low | Investigations and Diagnostic Tests | Respiratory System
| Lactate Dehydrogenase Increased | A laboratory test result indicating an increase in the lactate dehydrogenase concentration of a biological sample. | E2204 | 10023654 | Lactate dehydrogenase increased
| Cardiac Enzyme Elevation | Increase from the normal levels of cardiac enzymes. | E0603 | 10007548 | Cardiac enzymes increased | Heart | Investigations and Diagnostic Tests
| Cervical Changes | Changes such as abnormal, pathological, benign, malignant, etc. | E1404 | 10008351 | Cervix disorder | Reproductive System and Breast | Investigations and Diagnostic Tests
| Decreased Peak Expiratory Flow rate | Decrease in the maximum flow rate generated during a forceful exhalation, starting from full lung inflation. | E0713 | 10034195 | Peak expiratory flow rate decreased | Respiratory System | Investigations and Diagnostic Tests
| Increased Peak Expiratory Flow rate | An increase of the greatest rate of flow that can be achieved during forced expiration beginning with the lungs fully inflated. | E0714 | 10034196 | Peak expiratory flow rate increased | Respiratory System | Investigations and Diagnostic Tests
| Decreased Respiratory Rate | Abnormal decrease of rate of breathing. | E0715 | 10038710 | Respiratory rate decreased | Respiratory System | Investigations and Diagnostic Tests
| Increased Respiratory Rate | Abnormal increase of rate of breathing. | E0716 | 10038712 | Respiratory rate increased | Respiratory System | Investigations and Diagnostic Tests
| Forced Expiratory Volume Decreased | A decrease in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | E0719 | 10016987 | Forced expiratory volume decreased | Respiratory System | Investigations and Diagnostic Tests
| Forced Expiratory Volume Increased | An increase in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | E0720 | 10016989 | Forced expiratory volume increased | Respiratory System | Investigations and Diagnostic Tests
| High Pulmonary Arterial Wedge Pressure | The measurement of the mean left arterial pressure, as measured by a catheter introduced into the distal pulmonary artery, is high. | E0507 | 10037328 | Pulmonary arterial wedge pressure high | Vascular system | Investigations and Diagnostic Tests
| Low Pulmonary Arterial Wedge Pressure | The measurement of the mean left pulmonary arterial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low. | E0508 | 10037330 | Pulmonary arterial wedge pressure low | Vascular system | Investigations and Diagnostic Tests
| Hyperglycemia | Abnormally high level of glucose in the blood. | E1205 | 10020639 | Hyperglycemia | Metabolism, and Nutrition | Investigations and Diagnostic Tests
| Hypoglycemia | Abnormally low level of glucose in the blood. | E1206 | 10021005 | Hypoglycemia | Metabolism, and Nutrition | Investigations and Diagnostic Tests
| Low Apgar Score | Any Apgar score less than seven. This may be the initial or a subsequent apgar score. | E1510 | 10002944 | Apgar score low | Pregnancy, Childbirth and the Puerperium | Investigations and Diagnostic Tests
Generalized Disorders | Alteration in Body Temperature | A Change or alteration of Body Temperature to a temperature outside the range of normal body temperature. | E2301 | 10063488 | Body temperature fluctuation |  |  | Fever | Elevation of body temperature above normal due to the bodies own responses. | E230101 | 10016558 | Fever
[8 empty] | Hyperthermia | An abnormally high body temperature. Not fever. | E230102 | 10020843 | Hyperthermia
[8 empty] | Hypothermia | Abnormally low body temperature. | E230103 | 10021113 | Hypothermia
| Ambulation Difficulties | Any persistent problem with walking or gait. | E2302 | 10047810 | Walking difficulty | Generalized Disorders | Musculoskeletal System
| Capsular Contracture | The tightening of scar tissue that forms around the implant. | E2303 | 10080895 | Capsular contracture associated with implant
| Chills | Shivering or moderate tremors of the body often accompanied by a cold sensation. | E2304 | 10008531 | Chills
| Cyanosis | A bluish or purplish discoloration of the skin and mucous membranes resulting from a reduced amount of oxygenated hemoglobin in the blood. | E2305 | 10011703 | Cyanosis
| Decreased Appetite | Lack of desire for food. | E2306 | 10061428 | Decreased appetite | Generalized Disorders | Metabolism, and Nutrition
| Increased Appetite | Intense desire for food. | E2307 | 10021654 | Increased appetite | Generalized Disorders | Metabolism, and Nutrition
| Deformity/ Disfigurement | An alteration of appearance, shape, size or alignment. | E2308 | 10061619 | Deformity
| Deposits | Accumulation of extraneous inorganic matter in tissues or cavities. | E2309 | 10070692 | Device deposit issue |  |  | Calcium Deposits/Calcification | Deposition of calcium salt in the tissues. | E230901 | 10006958 | Calcium deposits
| Diaphoresis | Profuse sweating. | E2310 | 10012703 | Diaphoresis
| Discomfort | A feeling of mental or physical uneasiness, pain, or distress. | E2311 | 10013082 | Discomfort
| Fatigue | A state of generalized weakness with a pronounced inability to summon sufficient energy to accomplish daily activities. | E2312 | 10016256 | Fatigue |  |  | Extreme Exhaustion | Extreme fatique; the inability to repond to stimuli. | E231201 | 10015667 | Exhaustion
| Fibrosis | The formation of fibrous tissue. | E2313 | 10016642 | Fibrosis |  |  | Adhesion(s) | A fibrous band of tissue that connects normally separate body regions. | E2101 | 10059837 | Adhesion | Procedural Complications | Generalized Disorders
| Fistula | Abnormal epithelial-lined communication between two anatomical structures. | E2314 | 10016717 | Fistula | Generalized Disorders | Musculoskeletal System
| Fluid Discharge | Flow or leakage of fluid other than blood, from an orifice, wound or tissue. | E2315 | 10080911 | Fluid discharge |  |  | Purulent Discharge | A flow or draining of pus or exudate from an oriface or wound. | E231501 | 10037569 | Purulent discharge
| Foreign Body Reaction | An inflammatory reaction evoked by the presence of an exogenous material in the tissues. | E2316 | 10068159 | Foreign body reaction
| Granuloma | An inflammatory reaction usually caused by infectious organisms, foreign bodies, or cholesterol deposits. It is characterized by the presence of epithelioid histiocytes and chronic inflammation. | E2317 | 10018691 | Granuloma | Generalized Disorders | Skin and Subcutaneous Tissue
| Hair Loss | Loss of hair usually from the scalp that may result in bald spots or spread to the entire scalp or epidermis. | E2318 | 10019045 | Hair loss
| Hernia | The protrusion of part of an organ or fibroadipose tissue through an opening or weakened area. | E2319 | 10019909 | Hernia
| High Blood Pressure/ Hypertension | Blood pressure that is abnormally high. | E2320 | 10020772 | Hypertension
| Low Blood Pressure/ Hypotension | Blood pressure that is abnormally low. | E2321 | 10021097 | Hypotension
| Hyperemia | The presence of an increased amount of blood in a part or organ; engorgement. | E2322 | 10020619 | Hyperemia
| Hyperplasia | An abnormal increase in the number of cells in an organ or a tissue with consequent enlargement. | E2323 | 10020718 | Hyperplasia
| Incontinence | Involuntary passage of feces and/or urine from the body. | E2324 | 10021639 | Incontinence |  |  | Fecal Incontinence | Inability to control voluntary passage of feces from the body. | E232401 | 10016296 | Fecal incontinence
[8 empty] | Urinary Incontinence | Inability to control voluntary passage of urine from the body. | E232402 | 10046543 | Urinary incontinence
| Infiltration into Tissue | Permeation of foreign substances into tissues or cells. | E2325 | 10080901 | Tissue infiltration
| Inflammation | A localized physical condition in which part of the body becomes reddened, swollen, or hot and often painful. | E2326 | 10061218 | Inflammation
| Necrosis | The pathologic death of cells or of a portion of tissue or organ resulting from irreversible damage. | E2327 | 10028851 | Necrosis
| Obstruction/Occlusion | Blockage or closing of the normal flow of the contents of an anatomical passageway. | E2328 | 10061876 | Obstruction
| Organ Dehiscence | A splitting open or rupture of an internal organ with the exposure or discharge of its content. | E2329 | 10080904 | Dehiscence
| Pain | An unpleasant sensory and emotional experience that is associated with actual or potential tissue damage or described in such terms. | E2330 | 10033371 | Pain |  |  | Chest Pain | Pain in the chest. | E233001 | 10008479 | Chest pain
[8 empty] | Abdominal Pain | Painful sensation in the abdominal region. | E1002 | 10000081 | Abdominal pain | Gastrointestinal System | Generalized Disorders
[8 empty] | External Prosthetic Device Pain | Pain that is associated with an external ariticial replacement device. | E2106 | 10059057 | Medical device pain | Procedural Complications | Generalized Disorders
[8 empty] | Eye Pain | Painful sensation in the eye. | E0820 | 10015958 | Eye pain | Eye | Generalized Disorders
[8 empty] | Headache | Pain in various parts of the head, not confined to the area of distribution of any nerve. | E0116 | 10019211 | Headache | Nervous System | Generalized Disorders
[8 empty] | Inadequate Pain Relief | Continuing discomfort due to a lack or failure of a medical approach to treat pain. | E2110 | 10079788 | Inadequate pain relief | Procedural Complications | Generalized Disorders
[8 empty] | Intraoperative Pain | Pain occurring during a surgical operation. | E2112 | 10064882 | Procedural pain | Procedural Complications | Generalized Disorders
[8 empty] | Implant Pain | Pain localized to the site of the implanted device. | E2109 | 10063782 | Implant site pain | Procedural Complications | Generalized Disorders
[8 empty] | Myalgia | Pain sensation originating from a muscle or group of muscles. | E1622 | 10028411 | Myalgia | Musculoskeletal System | Generalized Disorders
[8 empty] | Neck Pain | Pain in the neck area. | E1623 | 10028836 | Neck pain | Musculoskeletal System | Generalized Disorders
| Pallor | Unhealthy, pale or white appearance of the face. | E2331 | 10033546 | Pallor
| Physical Asymmetry | Lack or absence of balanced proportions between parts or organs of the body. | E2332 | - | - | Generalized Disorders | Procedural Complications
| Prolapse | A condition in which an organ or body tissue drops or bulges out of place. | E2333 | 10076708 | Prolapse
| Ptosis | The drooping or downward displacement of a body part. | E2334 | 10037272 | Ptosis
| Reaction to Medicinal Component of Device | Any local or systemic adverse reaction to the medicinal component of a combination product. | E2335 | 10067484 | Adverse reaction
| Shock | A life-threatening condition that requires immediate medical intervention.It is characterized by reduced blood flow that may result in damage of multiple organs. | E2336 | 10040560 | Shock |  |  | Cardiogenic Shock | Shock resulting from primary failure of the heart in its pumping function, as in myocardial infarction, severe cardiomyopathy, or mechanical obstruction or compression of the heart. | E233601 | 10007625 | Cardiogenic shock
[8 empty] | Hypoglycemic Shock | A hypoglycemic reaction to overdosage of insulin, a skipped meal, or strenuous exercise in an insulin-dependent diabetic. | E233602 | 10060493 | Hypoglycemic shock
[8 empty] | Hypovolemic Shock | Shock resulting from insufficient blood volume for the maintenance of adequate cardiac output, blood pressure and tissue perfusion. | E233603 | 10021140 | Hypovolemic shock
[8 empty] | Neurogenic Shock | Shock resulting from neurogenic vasodilation, which can be produced by cerebral trauma or hemorrhage, spinal cord injury, deep general or spinal anesthesia, or toxic central nervous system depression. | E233604 | 10058119 | Neurogenic shock
[8 empty] | Septic Shock | A state of acute circulatory failure resulting from organ injury or damage in response to infection that results in dangerously low blood pressure. | E233605 | 10040070 | Septic shock
[8 empty] | Anaphylactic Shock | Progression of anaphylaxis into systemic shock, which is characterized by significant reduction in tissue perfusion which leads to hypoxia and hypotension, and can lead to death if untreated. | E040201 | 10002199 | Anaphylactic shock | Immune System | Generalized Disorders
| Stenosis | Narrowing or stricture of a vessel, duct or canal. | E2337 | 10076711 | Stenosis |  |  | Restenosis | The reoccurrence of stenosis after it has been treated. | E233701 | 10082493 | Restenosis
| Swelling/ Edema | Enlargement or expansion in size of a body part from injury or inflammation. | E2338 | 10014210 | Edema |  |  | Peripheral Edema | accumulation of fluid in tissues perfused by the peripheral vascular system. | E233801 | 10034570 | Peripheral edema
[8 empty] | Angioedema | Rapid swelling of the deep layers of the skin due to transient vascular leakage of serous fluid. | E1702 | 10002424 | Angioedema | Skin and Subcutaneous Tissue | Generalized Disorders
[8 empty] | Cerebral Edema | A swelling in the brain caused by the presence of excessive fluid. | E0103 | 10008107 | Cerebral edema | Nervous System | Generalized Disorders
[8 empty] | Corneal Edema | Hazy, swollen cornea. | E0807 | 10011007 | Corneal edema | Eye | Generalized Disorders
[8 empty] | Corneal Stromal Edema | Fluid from the aqueous humor enters the corneal stroma and causes swelling. | E080701 | 10054772 | Corneal stromal edema | Eye | Generalized Disorders
[8 empty] | Microcystic Edema | Corneal edema of inflammatory origin associated with contact lens wear and morphologically characterized by presence of epithelial microcysts. | E080702 | 10068529 | Corneal microcystic edema | Eye | Generalized Disorders
[8 empty] | Macular Edema | Accumulation of intraretinal fluid and protein in the macula, which may result in swelling and decreased central vision. | E0830 | 10054467 | Macular edema | Eye | Generalized Disorders
[8 empty] | Pulmonary Edema | Accumulation of fluid in the lung tissues causing disturbance of the gas exchange that may lead to respiratory failure. | E0736 | 10037375 | Pulmonary edema | Respiratory system | Generalized Disorders
| Ulcer | An open sore on an external or internal surface of the body caused by a break in the skin or mucous membrane that fails to heal. | E2339 | 10045285 | Ulcer
| Wound Dehiscence | A complication in which the wound splits or separates after being surgically closed. | E2340 | 10048031 | Wound dehiscence | Generalized Disorders | Procedural Complications
| Anaphylactoid | Resembling anaphylaxis | E040202 | 10002216 | Anaphylactoid reaction | Immune System | Generalized Disorders
| Burning Sensation | A sensation of stinging or heat, not necessarily accompanied by redness or physical signs of irritation. | E1705 | 10006784 | Burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders | Skin Burning Sensation | A sensation of dermal stinging or heat, not necessarily accompanied by redness or physical signs of skin irritation. | E170501 | 10054786 | Skin burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders
| Phosphene Visualization | An objective visual sensation that appears with the eyes closed and in the absence of visual light. | E083802 | 10034936 | Phosphenes | Eye | Generalized Disorders
Others | Insufficient Information | Health effect appears to have occurred but there is not yet enough information available to classify the clinical signs, symptoms and conditions. | E2401 | - | -
| Appropriate Term / Code Not Available | The clinical signs, symptoms and conditions are not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | E2402 | - | -
| No Clinical Signs, Symptoms or Conditions | No patient involvement or, no observable clinical symptoms or a change in symptoms is identified in the patient. | E2403 | - | -

## AnnexF Health Impact
Annex F. Health Impact
Device (bold): For the purpose of this Annex, a device means a medical device including accessories and components
Wherever appropriate "patient" should be taken to include user, operator or any other person affected by the incident.
LEVEL 1 |  |  | LEVEL 2 |  |  | LEVEL 3
Term | Definition | IMDRF Code | Term | Definition | IMDRF Code | Term | Definition | IMDRF Code
Change in Therapeutic Response | Change in response to treatment or cure of a disorder or disease. | F01 | Therapeutic Response Decreased | A reduction in the desirable and beneficial effects resulting from a medical treatment. | F0101
|  |  | Therapeutic Response Increased | An increase in the desirable and beneficial effects resulting from a medical treatment. | F0102
|  |  | Unexpected Therapeutic Effects | Unanticipated desirable and beneficial effects resulting from a medical treatment. | F0103
Death | The cessation of life. | F02 | Intrauterine Fetal Death | Death in utero; failure of the product of conception to show evidence of respiration, heart beat, or definite movement of a voluntary muscle after expulsion from the uterus, with no possibility of resuscitation. | F0201
Brain Death | The cessation of life due to permanent loss of brain function and brainstem for absence of responsiveness to external stimuli, absence of reflex and apnea. | F03
Delay to Diagnosis | Patient diagnosis was clinically significant delayed as a consequence of device performance. | F04
Delay to Treatment/ Therapy | Patient treatment was delayed as a consequence of device performance. | F05
Disruption of Subsequent Medical Procedure | Situation in which the use of the device impedes or affects a subsequent medical procedure or use of a medicine or device. The time elapsed between the use of the device and the medical procedure is not a factor. It is not necessary for the device to have broken or malfunctioned. | F06
Exacerbation of Existing Condition | Use of the device has lead to worsening of the existing disease or condition. | F07
Hospitalization or Prolonged Hospitalization | Hospitalization or prolonged hospitalization due to the health damage accompanying the use of device. | F08 | Intensive Care | Patient requires admission to or extension of stay in an intensive care unit. | F0801
Fetal Harm | Fetal distress, congenital physical or mental impairment or birth defect, excluding fetal death. | F09
Inadequate/Inappropriate Treatment or Diagnostic Exposure | Patient did not receive intended treatment from a device or received inadequate treatment as a consequence of device performance. This includes inadequate/inappropriate radiation dosage during procedures. | F10 | Absence of Treatment | Patient treatment was absent as a consequence of device performance. | F1001
|  |  | Incompatible Blood Transfusion | Receiving incorrect or incompatible whole blood or a blood component. | F1002
|  |  | Missed Dose | A dose of medicine was not taken or delivered at the prescribed dosing interval. | F1003
|  |  | Underdose | Exposure to, or intake of, an insufficient amount of a substance. | F1004 | Radiation Underdose | A lower than intended dose of radiation to a single field. | F100401
|  |  | Overdose | Exposure to, or intake of, an excessive amount of a substance. | F1005 | Radiation Overdose | A higher than intended cumulative dose of radiation to a single field. | F100501
Minor Injury/ Illness / Impairment | A mild injury, illness or impairment which can be treated with minimal or no intervention. | F11
Serious Injury/ Illness/ Impairment | A severe injury, illness or impairment which requires hospitalization or medical intervention. | F12 | Chronic Disease | Disease that persists over a long period of time, cannot be prevented by vaccines or cured by medication, nor do they just disappear. | F1201
|  |  | Disability | Any physical or mental impairment that interferes with an individual's ability to perform desired activities. | F1202
|  |  | Life Threatening Illness or Injury | Patient suffered an illness or injury which if not treated would be fatal. | F1203
|  |  | Permanent Impairment | Irreversible deterioration of the state of health. | F1204
|  |  | Temporary Impairment | Reversible deterioration of the state of health. | F1205
Misdiagnosis/ Misclassification | An incorrect diagnosis or classification of an illness or other problem. | F13
Prolonged Episode of Care | Clinically significant extended treatment time or procedureal time. Not to be used for prolonged hospitalization or prolonged surgery. | F14
Recognised Device or Procedural Complication | Patient experiencing a complication that is well documented in association with the device or procedure. | F15
Reduction in Life Expectancy | Decrease in lifetime/ lifespan of an individual. Decline in the normal duration or in average life. | F16
Sedation | Patient required sedation or additional sedation. | F17
Rehabilitation | Patient requires treatment to facilitate the process of recovery from injury, illness, or disease. | F18
Surgical Intervention | One or more surgical procedures was required, or an existing procedure changed. | F19 | Additional Surgery | Surgical intervention that was not planned and needed to be performed in addition to other interventions including surgical ones. | F1901
|  |  | Amputation | Patient required surgical removal of limb, extremities or other external body parts. | F1902
|  |  | Device Explantation | Patient required a corrective invasive procedure to remove a device, or parts/fragments of a device. | F1903
|  |  | Device Repositioning | Patient required a corrective invasive procedure in which a device was repositioned. This includes implantable, surgically invasive (penetration through the surface of the body) or invasive (penetration though a natural body orifice or permanent artificial opening, e.g. stoma) devices. | F1904
|  |  | Device Revision or Replacement | Patient required a corrective invasive procedure in which a device was replaced/revised. This includes implantable, surgically invasive (penetration through the surface of the body) or invasive (penetration though a natural body orifice or permanent artificial opening, e.g. stoma) devices. | F1905
|  |  | Modified Surgical Procedure | A change of surgical procedures from the original plan. | F1906
|  |  | More Complex Surgery | Surgery that is more complicated, more comprehensive or more extensive than planned. | F1907
|  |  | Prolonged surgery | Surgery that takes longer than anticipated. | F1908
|  |  | Surgical Procedure Delayed | Surgical procedure unable to be completed when originally scheduled and put off until some time in the future. | F1909
Serious Public Health Threat | Any event type, which results in imminent risk of death, serious injury, or serious illness to more than one individual, that requires prompt remedial action. | F20
Unexpected Deterioration | Unexpected clinical deterioration of patient. | F21
Unexpected Diagnostic Intervention | Diagnostic intervention needed to be completed that was not originally planned. | F22 | Biopsy | Patient required removal of tissue or fluid specimen to establish a diagnosis. | F2201
|  |  | Endoscopic Procedure | Patient required additional or unexpected endoscopic procedure. | F2202
|  |  | Imaging Required | Requirement for imaging that was not originally planned. | F2203
|  |  | IVD Testing | Pateint required additional or unexpected IVD testing. | F2204
Unexpected Medical Intervention | Patient required an unforeseen medical intervention, excluding surgery, which was not on the original treatment plan. | F23 | Additional Device Required | Use of an additional or alternative device require to achieve optimal outcome. | F2301
|  |  | Blood Transfusion | Patient required an infusion of whole blood or a blood component directly into the bloodstream. | F2302
|  |  | Medication Required | Patient required additional medication or additional dose of existing medication. | F2303
|  |  | Prophylactic Treatment | A medication or course of action used to prevent a possible disease or condition from occurring. | F2304
|  |  | Radiation Therapy | Patient required addtional radiation therapy. | F2305
|  |  | Resuscitation | Patient required resuscitation. | F2306
Insufficient Information | There is not yet enough information available to classify the health impact. | F24
Unanticipated Adverse Device Effect | Complication occuring in the patient that had not previously been identified or expected for this device. | F25
No Health Consequences or Impact | No apparent harm ocurred in relation to the adverse event. | F26
No Patient Involvement | No patient involvement when the adverse event occurred (for example happened during set-up or cleaning). | F27
Appropriate Term/Code Not Available | The Health Impact is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | F28